{"text_format":"1","workbook":"xlsx","n_sheets":7,"sheets":["Gam Cal","Gam  Plot","Displ Cal","Displ plot","P Wave Cal","Resistivity Cal","Resistivity plot"],"not_rendered":{"charts":6,"images":0,"macros":1,"external_links":0,"defined_names":7,"comments":0,"pivot_tables":0},"unopened":[{"sheet":"Gam  Plot","index":1,"reason":"not_a_worksheet"},{"sheet":"Displ plot","index":3,"reason":"not_a_worksheet"},{"sheet":"Resistivity plot","index":6,"reason":"not_a_worksheet"}],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"Gam Cal","index":0,"n_rows":14,"n_cols":5,"rows_kept":14,"cut":false,"hidden":false,"merged":[]}
["Gamma Calibration","User must enter values in Red"]
[]
["Internal Thickness (cm) =",6.6]
["Titanium Density (g/cc) =",4.5060000000000002]
["Water Density (g/cc) =",1]
[]
["Aluminium Thickness (cm)","Average Density (g/cc)","Av. Den. * Thickness","Intensity (cps)","Ln(cps)"]
[6,{"$f":"((A8/D)*Al_Dens)+(((D-A8)/D*W_Dens))","v":4.1872727272727275},{"$f":"B8*D","v":27.635999999999999},4055.9036144578313,{"$f":"LN(D8)","v":8.3079287812948301}]
[5,{"$f":"((A9/D)*Al_Dens)+(((D-A9)/D*W_Dens))","v":3.6560606060606058},{"$f":"B9*D","v":24.129999999999995},5090.5,{"$f":"LN(D9)","v":8.5351313365472503}]
[4,{"$f":"((A10/D)*Al_Dens)+(((D-A10)/D*W_Dens))","v":3.124848484848485},{"$f":"B10*D","v":20.623999999999999},6390.6265060240967,{"$f":"LN(D10)","v":8.7625875873267116}]
[3,{"$f":"((A11/D)*Al_Dens)+(((D-A11)/D*W_Dens))","v":2.5936363636363637},{"$f":"B11*D","v":17.117999999999999},7952.5714285714284,8.9812506054844299]
[2,{"$f":"((A12/D)*Al_Dens)+(((D-A12)/D*W_Dens))","v":2.0624242424242425},13.612,10243.763888888889,9.2344243983315035]
[1,{"$f":"((A13/D)*Al_Dens)+(((D-A13)/D*W_Dens))","v":1.5312121212121212},{"$f":"B13*D","v":10.106},13286.915492957747,9.4945350324414619]
[0,{"$f":"((A14/D)*Al_Dens)+(((D-A14)/D*W_Dens))","v":1},{"$f":"B14*D","v":6.6},14565.617977528091,9.5864190974346197]

{"sheet":"Displ Cal","index":2,"n_rows":14,"n_cols":2,"rows_kept":14,"cut":false,"hidden":false,"merged":[]}
["Core  Diameter Calibration","User must enter values in Red"]
[]
[" "," "]
["Volts (bits)","Displacement (mm)"]
[" "," "]
[1768.8928571428571,0]
[3325.5813953488373,-3.2]
[4246.159090909091,-5]
[6756,-10]
[9226.125,-15]
[" "," "]
[" "," "]
[]
["Scale on Displacement Channel","=slope"]

{"sheet":"P Wave Cal","index":4,"n_rows":33,"n_cols":10,"rows_kept":33,"cut":false,"hidden":false,"merged":[]}
["P-wave Calibration","User must enter values in Red"]
[]
["Temperature (deg C) =",22.6290625]
["Salinity (ppt) =",0]
["Depth (m) =",0]
["Velocity of sound in distilled water (m/s) =",{"$f":"1492.9+3*(T-10)-0.006*(T-10)^2-0.04*(T-18)^2+1.2*(S-35)-0.01*(T-18)*(S-35)+z/61","v":1488.5932712720705},null,null,null,null,null,"Velocity of sound in titanium (m/s) =",2730]
["Total Thickness (cm) =",7.12,null,null,null,null,null,"Total Thickness (cm) =",7.12]
["Total Liner Thickness (cm) =",0.52,null,null,null,null,null,"Total Liner Thickness (cm) =",0]
[]
["TT =",{"$f":"10000*(B7-B8)/B6","v":44.337161314453617},"(micro secs)",null,null,null,null,"TT =",{"$f":"10000*(I7-I8)/I6","v":26.08058608058608},"(micro secs)"]
["TOT =",70.349999999999952,"(micro secs)",null,null,null,null,"TOT =",47.4,"(micro secs)"]
["PTO =",{"$f":"B11-B10","v":26.012838685546335},"(micro secs)",null,null,null,null,"PTO =",{"$f":"I11-I10","v":21.319413919413918},"(micro secs)"]
[]
[null,"Salinity (ppt) =",0]
[null,"Depth (m) =",0]
[null,null,null,null,null,null,null,null,{"$f":"64/(33-I12)*1000","v":5479.1771199197183}]
[null,"Temp (deg C)","Velocity (m/s)",null,"Depth (m)","Velocity (m/s)"]
[null,0,{"$f":"1492.9+3*(B18-10)-0.006*(B18-10)^2-0.04*(B18-18)^2+1.2*($C$14-35)-0.01*(B18-18)*($C$14-35)+$C$15/61","v":1401.0400000000002},null,0,{"$f":"1492.9+3*(4-10)-0.006*(4-10)^2-0.04*(4-18)^2+1.2*(35-35)-0.01*(4-18)*(35-35)+E18/61","v":1466.8440000000003}]
[null,{"$f":"B18+2","v":2},1410.6760000000002,null,500,{"$f":"1492.9+3*(4-10)-0.006*(4-10)^2-0.04*(4-18)^2+1.2*(35-35)-0.01*(4-18)*(35-35)+E19/61","v":1475.0407213114756}]
[null,{"$f":"B19+2","v":4},1419.9440000000002,null,1000,1483.237442622951]
[null,6,1428.8440000000001,null,1500,1491.4341639344266]
[null,8,1437.3760000000002,null,2000,1499.6308852459019]
[null,10,1445.5400000000002,null,2500,1507.8276065573773]
[null,12,1453.3360000000002,null,3000,1516.0243278688527]
[null,14,1460.7639999999999,null,3500,1524.2210491803282]
[null,16,1467.8240000000001,null,4000,1532.4177704918036]
[null,{"$f":"B26+2","v":18},1474.5160000000001,null,4500,1540.614491803279]
[null,20,1480.8400000000001,null,5000,1548.8112131147543]
[null,22,1486.796,null,5500,1557.0079344262299]
[null,24,1492.384,null,6000,1565.2046557377053]
[null,26,1497.604]
[null,28,1502.4560000000001]
[null,30,1506.94]

{"sheet":"Resistivity Cal","index":5,"n_rows":9,"n_cols":11,"rows_kept":9,"cut":false,"hidden":false,"merged":[]}
["Saline","Resistivity","Temperature","Sensor","Corrected","Temp. Corrected"]
["Concentration","at 20degC","Gradient","Response","Sensor Res.","Resitivity"]
["(g/l)","(ohm-m)","(ohm-m/degC)","(mV)","(mV)","(ohm-m)",null,null,"Calibration Temperature",23,"degC"]
[35,0.20869182306555434,-4.0000000000000001E-3,434.81819999999999,{"$f":"D4-$D$9","v":409.53249},{"$f":"B4+C4*($J$3-20)","v":0.19669182306555433}]
[17.5,0.39019429810381373,-8.0000000000000002E-3,251.24,{"$f":"D5-$D$9","v":225.95429000000001},{"$f":"B5+C5*($J$3-20)","v":0.36619429810381371}]
[3.5,1.7298459416074625,-3.7999999999999999E-2,74.923079999999999,{"$f":"D6-$D$9","v":49.637369999999997},{"$f":"B6+C6*($J$3-20)","v":1.6158459416074624}]
[1.75,3.3209511658245541,-7.2999999999999995E-2,51.478259999999999,{"$f":"D7-$D$9","v":26.192549999999997},{"$f":"B7+C7*($J$3-20)","v":3.1019511658245542},null,null," Sediment Range"]
[0.35,15.48122979817003,-0.34100000000000003,31.409089999999999,{"$f":"D8-$D$9","v":6.1233799999999974},{"$f":"B8+C8*($J$3-20)","v":14.458229798170031}]
["Air","-","-",25.285710000000002,"-","-"]