## Caption SuppTable2
Supplemental Table 2.
Calibration Data for XRF1. (A) 10 KV, (B) 30 KV, (C) 50 KV. Each includes calibration data for USGS Standards on XRF1, Calibration Method #1: Sensitivity/First Approximation, Slope Corrections; Calibration Method #2: Traditional regression calibration; and standards data treated as unknowns, with calculated (first approximation), slope-corrected, and traditional regression analysis compared to reference values for each individual calibration count.

## STable2A-10 KV XRF1
USGS Standards | Calibration Data | MgO | Counts /Rh | Mg-Ka Area | Mg-Ka Stdev | Al2O3 | Counts /Rh | Al-Ka Area | Al-Ka Stdev | SiO2 | Counts /Rh | Si-Ka Area | Si-Ka Stdev | P2O5 | Counts /Rh | P -Ka Area | P -Ka Stdev | K2O | Counts /Rh | K -Ka Area | K -Ka Stdev | CaO | Counts /Rh | Ca-Ka Area | Ca-Ka Stdev | TiO2 | Counts /Rh | Ti-Ka Area | Ti-Ka Stdev | Cr ppm | Counts /Rh | Cr-Ka Area | Cr-Ka Stdev | MnO | Counts /Rh | Mn-Ka Area | Mn-Ka Stdev | FeO | Counts /Rh | Fe-Ka Area | Fe-Ka Stdev | Rh-La-Coh Area | Rh-La Coh Stdev | ratio rh
BIR1 |  | 9.68 | 0.088791 | 2808 | 81 | 15.35 | 0.8585 | 27150 | 183 | 47.77 | 5.5246 | 174716 | 423 | 0.046 | 0.011826 | 374 | 78 | 0.027 | 0.0014862 | 47 | 120 | 13.24 | 13.4994 | 426917 | 641 | 0.96 | 0.89723 | 28375 | 214 | 382 | 0.15219 | 4813 | 191 | 0.171 | 0.18786 | 5941 | 184 | 10.2517 | 11.2001 | 354204 | 602 | 31625 | 211 | 1.0026
BIR1 |  | 9.68 | 0.082859 | 2635 | 82 | 15.35 | 0.86343 | 27458 | 185 | 47.77 | 5.5242 | 175674 | 425 | 0.046 | 0.010974 | 349 | 78 | 0.027 | 0 |  | 121 | 13.24 | 13.346 | 424415 | 637 | 0.96 | 0.88485 | 28139 | 213 | 382 | 0.13927 | 4429 | 174 | 0.171 | 0.1871 | 5950 | 182 | 10.2517 | 11.0952 | 352838 | 598 | 31801 | 159 | 0.99705
BIR1 |  | 9.68 | 0.096583 | 3050 | 83 | 15.35 | 0.86988 | 27470 | 185 | 47.77 | 5.569 | 175865 | 425 | 0.046 | 0.013775 | 435 | 79 | 0.027 | 0 |  | 120 | 13.24 | 13.4736 | 425483 | 639 | 0.96 | 0.90259 | 28503 | 215 | 382 | 0.13674 | 4318 | 190 | 0.171 | 0.19405 | 6128 | 184 | 10.2517 | 11.14 | 351789 | 600 | 31579 | 210 | 1.0041
BIR1 |  | 9.68 | 0.095801 | 3037 | 82 | 15.35 | 0.87228 | 27652 | 184 | 47.77 | 5.5772 | 176803 | 426 | 0.046 | 0.010473 | 332 | 78 | 0.027 | 0.0079808 | 253 | 120 | 13.24 | 13.4437 | 426179 | 639 | 0.96 | 0.88568 | 28077 | 214 | 382 | 0.12807 | 4060 | 191 | 0.171 | 0.18249 | 5785 | 183 | 10.2517 | 11.1247 | 352664 | 600 | 31701 | 211 | 1.0002
BIR1 |  | 9.68 | 0.090929 | 2866 | 82 | 15.35 | 0.87297 | 27515 | 185 | 47.77 | 5.6067 | 176717 | 426 | 0.046 | 0.013294 | 419 | 80 | 0.027 | 0.0011104 | 35 | 120 | 13.24 | 13.4825 | 424956 | 639 | 0.96 | 0.90225 | 28438 | 213 | 382 | 0.14874 | 4688 | 190 | 0.171 | 0.18503 | 5832 | 184 | 10.2517 | 11.195 | 352854 | 601 | 31519 | 211 | 1.006
BHVO1 |  | 7.23 | 0.054737 | 1746 | 75 | 13.8 | 0.73014 | 23290 | 175 | 49.94 | 5.8926 | 187962 | 440 | 0.273 | 0.048812 | 1557 | 90 | 0.52 | 0.38711 | 12348 | 178 | 11.4 | 11.3152 | 360932 | 596 | 2.71 | 2.6583 | 84793 | 318 | 289 | 0.10107 | 3224 | 171 | 0.168 | 0.18462 | 5889 | 184 | 11.1176 | 11.8113 | 376757 | 616 | 31898 | 161 | 0.99401
BHVO1 |  | 7.23 | 0.057124 | 1827 | 77 | 13.8 | 0.7423 | 23741 | 177 | 49.94 | 6.0219 | 192598 | 445 | 0.273 | 0.055498 | 1775 | 91 | 0.52 | 0.38152 | 12202 | 178 | 11.4 | 11.3744 | 363788 | 598 | 2.71 | 2.6673 | 85309 | 320 | 289 | 0.086077 | 2753 | 172 | 0.168 | 0.17834 | 5704 | 184 | 11.1176 | 11.8664 | 379523 | 619 | 31983 | 161 | 0.99137
BHVO1 |  | 7.23 | 0.057069 | 1841 | 76 | 13.8 | 0.72829 | 23494 | 176 | 49.94 | 5.9854 | 193084 | 446 | 0.273 | 0.049351 | 1592 | 91 | 0.52 | 0.38023 | 12266 | 179 | 11.4 | 11.34 | 365818 | 602 | 2.71 | 2.639 | 85133 | 320 | 289 | 0.092098 | 2971 | 187 | 0.168 | 0.17112 | 5520 | 186 | 11.1176 | 11.7467 | 378937 | 622 | 32259 | 213 | 0.98289
BHVO1 |  | 7.23 | 0.063591 | 2048 | 77 | 13.8 | 0.74092 | 23862 | 176 | 49.94 | 5.9722 | 192340 | 445 | 0.273 | 0.050581 | 1629 | 91 | 0.52 | 0.38543 | 12413 | 179 | 11.4 | 11.2852 | 363452 | 598 | 2.71 | 2.6553 | 85516 | 319 | 289 | 0.10181 | 3279 | 171 | 0.168 | 0.1831 | 5897 | 183 | 11.1176 | 11.8207 | 380699 | 620 | 32206 | 162 | 0.98451
BHVO1 |  | 7.23 | 0.057986 | 1848 | 76 | 13.8 | 0.75253 | 23983 | 177 | 49.94 | 6.0688 | 193413 | 446 | 0.273 | 0.052024 | 1658 | 90 | 0.52 | 0.38553 | 12287 | 178 | 11.4 | 11.4367 | 364487 | 601 | 2.71 | 2.6485 | 84407 | 320 | 289 | 0.098651 | 3144 | 187 | 0.168 | 0.18036 | 5748 | 184 | 11.1176 | 11.9422 | 380598 | 622 | 31870 | 214 | 0.99489
BCR1 |  | 3.48 | 0.014514 | 449 | 60 | 13.65 | 0.75319 | 23300 | 173 | 54.11 | 6.7573 | 209036 | 462 | 0.36 | 0.070729 | 2188 | 89 | 1.69 | 1.4115 | 43665 | 260 | 6.95 | 6.8826 | 212914 | 476 | 2.24 | 2.3676 | 73243 | 300 | 16 | 0.032487 | 1005 | 163 | 0.18 | 0.21503 | 6652 | 188 | 12.1105 | 13.6972 | 423723 | 651 | 30935 | 159 | 1.025
BCR1 |  | 3.48 | 0.016754 | 528 | 62 | 13.65 | 0.76953 | 24251 | 177 | 54.11 | 6.8776 | 216742 | 471 | 0.36 | 0.076315 | 2405 | 92 | 1.69 | 1.4191 | 44723 | 263 | 6.95 | 6.8664 | 216388 | 482 | 2.24 | 2.333 | 73523 | 302 | 16 | 0.033699 | 1062 | 174 | 0.18 | 0.21619 | 6813 | 190 | 12.1105 | 13.618 | 429159 | 658 | 31514 | 211 | 1.0061
BCR1 |  | 3.48 | 0.013605 | 430 | 62 | 13.65 | 0.76925 | 24313 | 177 | 54.11 | 6.8605 | 216833 | 470 | 0.36 | 0.078783 | 2490 | 92 | 1.69 | 1.4038 | 44369 | 263 | 6.95 | 6.8902 | 217773 | 482 | 2.24 | 2.3478 | 74203 | 302 | 16 | 0.036259 | 1146 | 165 | 0.18 | 0.2236 | 7067 | 190 | 12.1105 | 13.5925 | 429606 | 656 | 31606 | 161 | 1.0032
BCR1 |  | 3.48 | 0.015888 | 505 | 61 | 13.65 | 0.77039 | 24486 | 177 | 54.11 | 6.8669 | 218259 | 472 | 0.36 | 0.07831 | 2489 | 93 | 1.69 | 1.407 | 44719 | 263 | 6.95 | 6.8368 | 217302 | 481 | 2.24 | 2.329 | 74025 | 302 | 16 | 0.032847 | 1044 | 164 | 0.18 | 0.20406 | 6486 | 188 | 12.1105 | 13.5003 | 429092 | 655 | 31784 | 160 | 0.99758
BCR1 |  | 3.48 | 0.015195 | 476 | 62 | 13.65 | 0.76984 | 24116 | 176 | 54.11 | 6.9704 | 218355 | 472 | 0.36 | 0.076358 | 2392 | 93 | 1.69 | 1.4312 | 44835 | 263 | 6.95 | 6.9466 | 217609 | 483 | 2.24 | 2.3518 | 73672 | 303 | 16 | 0.031699 | 993 | 174 | 0.18 | 0.22196 | 6953 | 192 | 12.1105 | 13.647 | 427505 | 658 | 31326 | 211 | 1.0122
Method #1: Sensitivity Counts/Unit Concentation
Element |  | MgO |  |  |  | Al2O3 |  |  |  | SiO2 |  |  |  | P2O5 |  |  |  | K2O |  |  |  | CaO |  |  |  | TiO2 |  |  |  | Cr ppm |  |  |  | MnO |  |  |  | FeO |  |  |  | Rh-La Coh
|  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD
BIR1 |  | 297.438 | 2879.2 | 2.861% |  | 1788.2085 | 27449 | 0.6724% |  | 3683.3787 | 175955 | 0.2411% |  | 8300 | 381.8 | 19.09% |  | 4135.8025 | 111.6667 | 342.9% |  | 32144.2598 | 425590 | 0.1504% |  | 29485.8333 | 28306.4 | 0.749% |  | 11.6796 | 4461.6 | 4.053% |  | 34661.9883 | 5927.2 | 3.155% |  | 34420.6742 | 352869.8 | 0.1703% |  | 31645 | 0.6694%
BHVO1 |  | 257.538 | 1862 | 4.092% |  | 1715.5072 | 23674 | 0.7443% |  | 3842.1986 | 191879.4 | 0.2316% |  | 6015.3846 | 1642.2 | 5.517% |  | 23660 | 12303.2 | 1.45% |  | 31903.1053 | 363695.4 | 0.1647% |  | 31376.9742 | 85031.6 | 0.3756% |  | 10.6374 | 3074.2 | 5.777% |  | 34235.7143 | 5751.6 | 3.203% |  | 34117.4238 | 379302.8 | 0.1634% |  | 32043.2 | 0.5686%
BCR1 |  | 137.2414 | 477.6 | 12.86% |  | 1765.0696 | 24093.2 | 0.7305% |  | 3989.0039 | 215845 | 0.2175% |  | 6646.6667 | 2392.8 | 3.837% |  | 26308.9941 | 44462.2 | 0.5902% |  | 31136.2878 | 216397.2 | 0.2222% |  | 32916.6071 | 73733.2 | 0.4093% |  | 65.625 | 1050 | 16% |  | 37745.5556 | 6794.2 | 2.791% |  | 35326.2654 | 427817 | 0.1532% |  | 31433 | 0.5739%
Total Counts | Avg | 277.488 |  |  |  | 1756.2618 | 25072.0667 |  |  | 3838.1937 |  |  |  | 6331.0256 |  |  |  | 24984.497 |  |  |  | 31727.8843 |  |  |  | 31259.8049 |  |  |  | 11.1585 |  |  |  | 35547.7527 |  |  |  | 34621.4545 |  |  |  | 31707.0667
/wt% or ppm | stdev | 28.2135 |  |  |  | 37.1423 |  |  |  | 152.852 |  |  |  | 446.3838 |  |  |  | 1873.1217 |  |  |  | 526.3352 |  |  |  | 1718.3855 |  |  |  | 0.73695 |  |  |  | 1915.2494 |  |  |  | 628.9348 |  |  |  | 309.7987
| %std | 10.17% |  |  |  | 2.115% |  |  |  | 3.982% |  |  |  | 7.051% |  |  |  | 7.497% |  |  |  | 1.659% |  |  |  | 5.497% |  |  |  | 6.604% |  |  |  | 5.388% |  |  |  | 1.817% |  |  |  | 0.9771%
Method #2:
Slope Correction |  | MgO |  |  |  | Al2O3 |  |  |  | SiO2 |  |  |  | P2O5 |  |  |  | K2O |  |  |  | CaO |  |  |  | TiO2 |  |  |  | Cr ppm |  |  |  | MnO |  |  |  | FeO
Slope/ | Intercept | 0.70967 | 2.356 |  |  | 0.72695 | 3.8868 |  |  | 0.57975 | 21.2082 |  |  | 0.98774 | -0.0035629 |  |  | 0.92075 | 0.044329 |  |  | 0.96452 | 0.35012 |  |  | 0.94572 | 0.085387 |  |  | 1.195 | -77.0027 |  |  | 0.32541 | 0.1166 |  |  | 0.77997 | 2.4454
std err slope | std err intercept | 0.023361 | 0.16827 |  |  | 0.065048 | 0.93109 |  |  | 0.016417 | 0.83577 |  |  | 0.041517 | 0.011119 |  |  | 0.0077039 | 0.0082737 |  |  | 0.009288 | 0.10127 |  |  | 0.026021 | 0.055666 |  |  | 0.077262 | 22.0692 |  |  | 0.0306 | 0.0053233 |  |  | 0.047828 | 0.5363
r2 | std err y est | 0.98611 | 0.32279 |  |  | 0.90573 | 0.25346 |  |  | 0.98968 | 0.28704 |  |  | 0.97755 | 0.021305 |  |  | 0.99909 | 0.02259 |  |  | 0.9988 | 0.098421 |  |  | 0.99025 | 0.078418 |  |  | 0.94846 | 37.8777 |  |  | 0.8969 | 0.0017587 |  |  | 0.9534 | 0.1761
F-stat | deg freedom | 922.8634 | 13 |  |  | 124.8945 | 13 |  |  | 1247.1197 | 13 |  |  | 566.0364 | 13 |  |  | 14284.4546 | 13 |  |  | 10783.9028 | 13 |  |  | 1320.9602 | 13 |  |  | 239.2379 | 13 |  |  | 113.0931 | 13 |  |  | 265.948 | 13
regression sum of squares. | residual sum squares | 96.1539 | 1.3545 |  |  | 8.0232 | 0.83512 |  |  | 102.7513 | 1.0711 |  |  | 0.25692 | 0.0059007 |  |  | 7.2892 | 0.0066338 |  |  | 104.4611 | 0.12593 |  |  | 8.1231 | 0.079942 |  |  | 343238.6826 | 18651.3174 |  |  | 0.00034979 | 4.0208e-05 |  |  | 8.2478 | 0.40317
Method #3: Traditional |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept
Rh-tube line Corrected Counts |  | 81.0896 | 2.356 |  |  | 13.1241 | 3.8868 |  |  | 4.7893 | 21.2082 |  |  | 4.9468 | -0.0035629 |  |  | 1.1685 | 0.044329 |  |  | 0.96389 | 0.35012 |  |  | 0.95925 | 0.085387 |  |  | 3395.7321 | -77.0027 |  |  | 0.29026 | 0.1166 |  |  | 0.71431 | 2.4454
|  | 2.6693 | 0.16827 |  |  | 1.1744 | 0.93109 |  |  | 0.13562 | 0.83577 |  |  | 0.20792 | 0.011119 |  |  | 0.0097768 | 0.0082737 |  |  | 0.0092819 | 0.10127 |  |  | 0.026393 | 0.055666 |  |  | 219.5424 | 22.0692 |  |  | 0.027294 | 0.0053233 |  |  | 0.043802 | 0.5363
|  | 0.98611 | 0.32279 |  |  | 0.90573 | 0.25346 |  |  | 0.98968 | 0.28704 |  |  | 0.97755 | 0.021305 |  |  | 0.99909 | 0.02259 |  |  | 0.9988 | 0.098421 |  |  | 0.99025 | 0.078418 |  |  | 0.94846 | 37.8777 |  |  | 0.8969 | 0.0017587 |  |  | 0.9534 | 0.1761
|  | 922.8634 | 13 |  |  | 124.8945 | 13 |  |  | 1247.1197 | 13 |  |  | 566.0364 | 13 |  |  | 14284.4546 | 13 |  |  | 10783.9028 | 13 |  |  | 1320.9602 | 13 |  |  | 239.2379 | 13 |  |  | 113.0931 | 13 |  |  | 265.948 | 13
|  | 96.1539 | 1.3545 |  |  | 8.0232 | 0.83512 |  |  | 102.7513 | 1.0711 |  |  | 0.25692 | 0.0059007 |  |  | 7.2892 | 0.0066338 |  |  | 104.4611 | 0.12593 |  |  | 8.1231 | 0.079942 |  |  | 343238.6826 | 18651.3174 |  |  | 0.00034979 | 4.0208e-05 |  |  | 8.2478 | 0.40317
133.8192
Standards as Unknowns |  | MgO | Mg calc | Mg Slope | MgO trad | Al2O3 | Al calc | Al Slope | Al2O3 Trad | SiO2 | SiO2 calc | SiO2 slope | SiO2 trad | P2O5 | P2O5 calc | P2O5 slope | P2O5 trad | K2O | K2O calc | K2O slope | K2O trad | CaO | CaO calc | CaO slope | CaO trad | TiO2 | TiO2 calc | TiO2 slope | TiO2 trad | Cr ppm | Cr calc | Cr slope | Cr trad | MnO | MnO calc | MnO slope | MnO trad | FeO | FeO calc | FeO* slope | FeO* trad | Rh-La Coh
BIR1 | XRF1 | 9.68 | 10.1456 | 9.556 | 9.556 | 15.35 | 15.4991 | 15.1539 | 15.1539 | 47.77 | 45.6385 | 47.6673 | 47.6673 | 0.046 | 0.059227 | 0.054939 | 0.054939 | 0.027 | 0.001886 | 0.046066 | 0.046066 | 13.24 | 13.4905 | 13.362 | 13.362 | 0.96 | 0.91007 | 0.94606 | 0.94606 | 382 | 432.4506 | 439.7928 | 439.7928 | 0.171 | 0.16756 | 0.17112 | 0.17112 | 10.2517 | 10.2573 | 10.4458 | 10.4458
BIR1 | XRF1 | 9.68 | 9.4679 | 9.0751 | 9.0751 | 15.35 | 15.5882 | 15.2186 | 15.2186 | 47.77 | 45.6348 | 47.6651 | 47.6651 | 0.046 | 0.054963 | 0.050726 | 0.050726 | 0.027 | 0 | 0.044329 | 0.044329 | 13.24 | 13.3372 | 13.2141 | 13.2141 | 0.96 | 0.89751 | 0.93417 | 0.93417 | 382 | 395.7456 | 395.9289 | 395.9289 | 0.171 | 0.16689 | 0.1709 | 0.1709 | 10.2517 | 10.1612 | 10.3708 | 10.3708
BIR1 | XRF1 | 9.68 | 11.036 | 10.1879 | 10.1879 | 15.35 | 15.7046 | 15.3033 | 15.3033 | 47.77 | 46.0055 | 47.8801 | 47.8801 | 0.046 | 0.068988 | 0.06458 | 0.06458 | 0.027 | 0 | 0.044329 | 0.044329 | 13.24 | 13.4648 | 13.3371 | 13.3371 | 0.96 | 0.91551 | 0.9512 | 0.9512 | 382 | 388.5398 | 387.3176 | 387.3176 | 0.171 | 0.17309 | 0.17292 | 0.17292 | 10.2517 | 10.2022 | 10.4028 | 10.4028
BIR1 | XRF1 | 9.68 | 10.9467 | 10.1246 | 10.1246 | 15.35 | 15.7478 | 15.3347 | 15.3347 | 47.77 | 46.0729 | 47.9191 | 47.9191 | 0.046 | 0.05245 | 0.048245 | 0.048245 | 0.027 | 0.010128 | 0.053655 | 0.053655 | 13.24 | 13.4349 | 13.3083 | 13.3083 | 0.96 | 0.89835 | 0.93497 | 0.93497 | 382 | 363.9186 | 357.8944 | 357.8944 | 0.171 | 0.16277 | 0.16956 | 0.16956 | 10.2517 | 10.1882 | 10.3919 | 10.3919
BIR1 | XRF1 | 9.68 | 10.39 | 9.7295 | 9.7295 | 15.35 | 15.7603 | 15.3438 | 15.3438 | 47.77 | 46.3164 | 48.0603 | 48.0603 | 0.046 | 0.066577 | 0.062198 | 0.062198 | 0.027 | 0.0014092 | 0.045626 | 0.045626 | 13.24 | 13.4737 | 13.3457 | 13.3457 | 0.96 | 0.91516 | 0.95087 | 0.95087 | 382 | 422.6359 | 428.0638 | 428.0638 | 0.171 | 0.16504 | 0.1703 | 0.1703 | 10.2517 | 10.2526 | 10.4421 | 10.4421
BHVO1 | XRF1 | 7.23 | 6.2545 | 6.7947 | 6.7947 | 13.8 | 13.1817 | 13.4693 | 13.4693 | 49.94 | 48.6783 | 49.4296 | 49.4296 | 0.273 | 0.24446 | 0.2379 | 0.2379 | 0.52 | 0.49127 | 0.49666 | 0.49666 | 11.4 | 11.3078 | 11.2567 | 11.2567 | 2.71 | 2.6963 | 2.6353 | 2.6353 | 289 | 287.1989 | 266.2113 | 266.2113 | 0.168 | 0.16467 | 0.17018 | 0.17018 | 11.1176 | 10.817 | 10.8824 | 10.8824
BHVO1 | XRF1 | 7.23 | 6.5273 | 6.9882 | 6.9882 | 13.8 | 13.4013 | 13.6289 | 13.6289 | 49.94 | 49.7464 | 50.0488 | 50.0488 | 0.273 | 0.27795 | 0.27098 | 0.27098 | 0.52 | 0.48417 | 0.49013 | 0.49013 | 11.4 | 11.367 | 11.3138 | 11.3138 | 2.71 | 2.7055 | 2.644 | 2.644 | 289 | 244.5897 | 215.2916 | 215.2916 | 0.168 | 0.15908 | 0.16836 | 0.16836 | 11.1176 | 10.8675 | 10.9217 | 10.9217
BHVO1 | XRF1 | 7.23 | 6.521 | 6.9838 | 6.9838 | 13.8 | 13.1484 | 13.4451 | 13.4451 | 49.94 | 49.4452 | 49.8742 | 49.8742 | 0.273 | 0.24716 | 0.24057 | 0.24057 | 0.52 | 0.48254 | 0.48863 | 0.48863 | 11.4 | 11.3326 | 11.2806 | 11.2806 | 2.71 | 2.6768 | 2.6169 | 2.6169 | 289 | 261.6995 | 235.7385 | 235.7385 | 0.168 | 0.15263 | 0.16626 | 0.16626 | 11.1176 | 10.7579 | 10.8362 | 10.8362
BHVO1 | XRF1 | 7.23 | 7.2662 | 7.5126 | 7.5126 | 13.8 | 13.3763 | 13.6108 | 13.6108 | 49.94 | 49.3358 | 49.8108 | 49.8108 | 0.273 | 0.25332 | 0.24665 | 0.24665 | 0.52 | 0.48913 | 0.4947 | 0.4947 | 11.4 | 11.2778 | 11.2278 | 11.2278 | 2.71 | 2.6933 | 2.6325 | 2.6325 | 289 | 289.3049 | 268.7281 | 268.7281 | 0.168 | 0.16332 | 0.16974 | 0.16974 | 11.1176 | 10.8257 | 10.8891 | 10.8891
BHVO1 | XRF1 | 7.23 | 6.6257 | 7.0581 | 7.0581 | 13.8 | 13.5859 | 13.7631 | 13.7631 | 49.94 | 50.134 | 50.2736 | 50.2736 | 0.273 | 0.26055 | 0.25379 | 0.25379 | 0.52 | 0.48927 | 0.49483 | 0.49483 | 11.4 | 11.4292 | 11.3738 | 11.3738 | 2.71 | 2.6864 | 2.6259 | 2.6259 | 289 | 280.3184 | 257.9889 | 257.9889 | 0.168 | 0.16087 | 0.16895 | 0.16895 | 11.1176 | 10.9369 | 10.9759 | 10.9759
BCR1 | XRF1 | 3.48 | 1.6585 | 3.533 | 3.533 | 13.65 | 13.5979 | 13.7718 | 13.7718 | 54.11 | 55.8213 | 53.5708 | 53.5708 | 0.36 | 0.35422 | 0.34632 | 0.34632 | 1.69 | 1.7913 | 1.6937 | 1.6937 | 6.95 | 6.8781 | 6.9842 | 6.9842 | 2.24 | 2.4015 | 2.3565 | 2.3565 | 16 | 92.3139 | 33.316 | 33.316 | 0.18 | 0.1918 | 0.17901 | 0.17901 | 12.1105 | 12.5442 | 12.2295 | 12.2295
BCR1 | XRF1 | 3.48 | 1.9144 | 3.7147 | 3.7147 | 13.65 | 13.8929 | 13.9863 | 13.9863 | 54.11 | 56.8157 | 54.1473 | 54.1473 | 0.36 | 0.3822 | 0.37396 | 0.37396 | 1.69 | 1.801 | 1.7026 | 1.7026 | 6.95 | 6.8619 | 6.9686 | 6.9686 | 2.24 | 2.3664 | 2.3233 | 2.3233 | 16 | 95.7574 | 37.4311 | 37.4311 | 0.18 | 0.19283 | 0.17935 | 0.17935 | 12.1105 | 12.4717 | 12.173 | 12.173
BCR1 | XRF1 | 3.48 | 1.5546 | 3.4593 | 3.4593 | 13.65 | 13.8879 | 13.9826 | 13.9826 | 54.11 | 56.6742 | 54.0652 | 54.0652 | 0.36 | 0.39456 | 0.38616 | 0.38616 | 1.69 | 1.7815 | 1.6847 | 1.6847 | 6.95 | 6.8857 | 6.9915 | 6.9915 | 2.24 | 2.3813 | 2.3375 | 2.3375 | 16 | 103.0306 | 46.1229 | 46.1229 | 0.18 | 0.19944 | 0.1815 | 0.1815 | 12.1105 | 12.4483 | 12.1547 | 12.1547
BCR1 | XRF1 | 3.48 | 1.8155 | 3.6444 | 3.6444 | 13.65 | 13.9084 | 13.9975 | 13.9975 | 54.11 | 56.7274 | 54.0961 | 54.0961 | 0.36 | 0.39219 | 0.38382 | 0.38382 | 1.69 | 1.7855 | 1.6884 | 1.6884 | 6.95 | 6.8324 | 6.94 | 6.94 | 2.24 | 2.3623 | 2.3195 | 2.3195 | 16 | 93.3347 | 34.5359 | 34.5359 | 0.18 | 0.18202 | 0.17583 | 0.17583 | 12.1105 | 12.3638 | 12.0888 | 12.0888
BCR1 | XRF1 | 3.48 | 1.7363 | 3.5882 | 3.5882 | 13.65 | 13.8985 | 13.9903 | 13.9903 | 54.11 | 57.5821 | 54.5916 | 54.5916 | 0.36 | 0.38242 | 0.37417 | 0.37417 | 1.69 | 1.8163 | 1.7167 | 1.7167 | 6.95 | 6.942 | 7.0458 | 7.0458 | 2.24 | 2.3854 | 2.3413 | 2.3413 | 16 | 90.0732 | 30.6383 | 30.6383 | 0.18 | 0.19798 | 0.18102 | 0.18102 | 12.1105 | 12.4982 | 12.1936 | 12.1936
|  | MgO | Mg calc | Mg Slope | MgO trad | Al2O3 | Al calc | Al Slope | Al2O3 Trad | SiO2 | SiO2 calc | SiO2 slope | SiO2 trad | P2O5 | P2O5 calc | P2O5 slope | P2O5 trad | K2O | K2O calc | K2O slope | K2O trad | CaO | CaO calc | CaO slope | CaO trad | TiO2 | TiO2 calc | TiO2 slope | TiO2 trad | Cr ppm | Cr calc | Cr slope | Cr trad | MnO | MnO calc | MnO slope | MnO trad | FeO | FeO calc | FeO* slope | FeO* trad
BIR-1 |  | 9.68 | 10.3972 | 9.7346 | 9.7346 | 15.35 | 15.66 | 15.2709 | 15.2709 | 47.77 | 45.9336 | 47.8384 | 47.8384 | 0.046 | 0.060441 | 0.056137 | 0.056137 | 0.027 | 0.0026847 | 0.046801 | 0.046801 | 13.24 | 13.4402 | 13.3134 | 13.3134 | 0.96 | 0.90732 | 0.94345 | 0.94345 | 382 | 400.6581 | 401.7995 | 401.7995 | 0.171 | 0.16707 | 0.17096 | 0.17096 | 10.2517 | 10.2123 | 10.4107 | 10.4107
BHVO-1 |  | 7.23 | 6.6389 | 7.0675 | 7.0675 | 13.8 | 13.3387 | 13.5834 | 13.5834 | 49.94 | 49.468 | 49.8874 | 49.8874 | 0.273 | 0.25669 | 0.24998 | 0.24998 | 0.52 | 0.48728 | 0.49299 | 0.49299 | 11.4 | 11.3429 | 11.2905 | 11.2905 | 2.71 | 2.6916 | 2.6309 | 2.6309 | 289 | 272.6223 | 248.7917 | 248.7917 | 0.168 | 0.16011 | 0.1687 | 0.1687 | 11.1176 | 10.841 | 10.9011 | 10.9011
BCR-1 |  | 3.48 | 1.7358 | 3.5879 | 3.5879 | 13.65 | 13.8371 | 13.9457 | 13.9457 | 54.11 | 56.7241 | 54.0942 | 54.0942 | 0.36 | 0.38112 | 0.37289 | 0.37289 | 1.69 | 1.7951 | 1.6972 | 1.6972 | 6.95 | 6.88 | 6.986 | 6.986 | 2.24 | 2.3794 | 2.3356 | 2.3356 | 16 | 94.902 | 36.4088 | 36.4088 | 0.18 | 0.19281 | 0.17934 | 0.17934 | 12.1105 | 12.4652 | 12.1679 | 12.1679
PDF VERSIONS | [13 empty] | PDF VERSIONS | [9 empty] | PDF VERSIONS | [8 empty] | PDF VERSIONS |  |  |  |  | PDF VERSIONS | [5 empty] | PDF VERSIONS
#VALUE!

## STable2B-30 KV XRF1
USGS Standards | Calibration Data | K2O | counts /rh | K -Ka Area | K -Ka Stdev | CaO | counts /rh | Ca-Ka Area | Ca-Ka Stdev | TiO2 | counts /rh | Ti-Ka Area | Ti-Ka Stdev | MnO | counts /rh | Mn-Ka Area | Mn-Ka Stdev | FeO | counts /rh | Fe-Ka Area | Fe-Ka Stdev | Co ppm | counts/rh | Co-Ka Area | Co-Ka Stdev | Ni ppm | counts /compton | Ni-Ka Area | Ni-Ka Stdev | Cu ppm | counts /compton | Cu-Ka Area | Cu-Ka Stdev | Zn ppm | counts /compton | Zn-Ka Area | Zn-Ka Stdev | Rb ppm | counts /rh | Rb-Ka Area | Rb-Ka Stdev | Sr ppm | counts /rh | Sr-Ka Area | Sr-Ka Stdev | Y ppm | counts /rh | Y -Ka Area | Y -Ka Stdev | Zr ppm | counts /rh | Zr-Ka Area | Zr-Ka Stdev | Nb ppm | counts /rh | Nb-Ka Area | Nb-Ka Stdev | Rh-Ka-Coh Area | Rh-Ka-Coh Stdev | Rh-Ka-Inc Area | Rh-Ka-Inc Stdev | Ro co/inc | Rh avg/co/inc
BIR1 |  | 0.027 | 137.9023 | 114 | 23 | 13.24 | 28017.1424 | 23161 | 147 | 0.96 | 3604.8135 | 2980 | 61 | 0.171 | 3159.6553 | 2612 | 63 | 10.2517 | 274574.2859 | 226983 | 455 | 51.4 | 3523.7656 | 2913 | 97 | 166 | 5690.2827 | 4704 | 76 | 126 | 1331.8455 | 1101 | 7.3 | 71 | 1236.2817 | 1022 | 6.8 | 0.27 | 516.5286 | 427 | 59 | 108 | 5843.9107 | 4831 | 85 | 16 | 944.7514 | 781 | 69 | 22 | 1437.0867 | 1188 | 95 | 2 | 229.8371 | 190 | 109 | 24739 | 274 | 29926 | 457 | 0.82667 | 1.035
BIR1 |  | 0.027 | 128.7178 | 107 | 25 | 13.24 | 27597.3274 | 22941 | 147 | 0.96 | 3423.6517 | 2846 | 60 | 0.171 | 3160.2013 | 2627 | 63 | 10.2517 | 270202.6286 | 224613 | 454 | 51.4 | 3405.6072 | 2831 | 98 | 166 | 5599.8239 | 4655 | 77 | 126 | 1378.6033 | 1146 | 7.3 | 71 | 1283.5686 | 1067 | 7 | 0.27 | 452.3166 | 376 | 60 | 108 | 5686.4377 | 4727 | 85 | 16 | 968.3906 | 805 | 68 | 22 | 1313.6429 | 1092 | 94 | 2 | 222.5494 | 185 | 110 | 24043 | 274 | 28923 | 455 | 0.83128 | 1.0293
BIR1 |  | 0.027 | 120.1637 | 100 | 24 | 13.24 | 27780.645 | 23119 | 147 | 0.96 | 3657.7829 | 3044 | 60 | 0.171 | 2925.986 | 2435 | 63 | 10.2517 | 270677.1378 | 225257 | 454 | 51.4 | 3555.6438 | 2959 | 98 | 166 | 5670.5248 | 4719 | 76 | 126 | 1359.0514 | 1131 | 7.5 | 71 | 1149.9666 | 957 | 6.7 | 0.27 | 675.32 | 562 | 60 | 108 | 5966.1275 | 4965 | 86 | 16 | 939.6801 | 782 | 69 | 22 | 1250.9041 | 1041 | 94 | 2 | 68.4933 | 57 | 110 | 24802 | 275 | 29803 | 456 | 0.8322 | 1.0281
BIR1 |  | 0.027 | 173.4997 | 146 | 22 | 13.24 | 27597.1435 | 23223 | 147 | 0.96 | 3473.5586 | 2923 | 60 | 0.171 | 2945.9294 | 2479 | 64 | 10.2517 | 266405.1854 | 224180 | 453 | 51.4 | 3489.0072 | 2936 | 100 | 166 | 5241.8292 | 4411 | 76 | 126 | 1321.4496 | 1112 | 7.5 | 71 | 1136.0664 | 956 | 6.7 | 0.27 | 588.2352 | 495 | 60 | 108 | 5644.6812 | 4750 | 85 | 16 | 767.6766 | 646 | 68 | 22 | 1188.3539 | 1000 | 94 | 2 | 79.6197 | 67 | 110 | 24592 | 275 | 29224 | 454 | 0.8415 | 1.0168
BIR1 |  | 0.027 | 129.5813 | 109 | 25 | 13.24 | 27237.0465 | 22911 | 146 | 0.96 | 3604.5011 | 3032 | 60 | 0.171 | 3112.3298 | 2618 | 63 | 10.2517 | 267388.1215 | 224919 | 454 | 51.4 | 3407.157 | 2866 | 99 | 166 | 5517.3119 | 4641 | 77 | 126 | 1236.3724 | 1040 | 7.4 | 71 | 1147.2109 | 965 | 6.8 | 0.27 | 530.2136 | 446 | 60 | 108 | 5810.9503 | 4888 | 86 | 16 | 848.8172 | 714 | 69 | 22 | 1198.3302 | 1008 | 94 | 2 | 166.4347 | 140 | 110 | 24865 | 276 | 29560 | 463 | 0.84117 | 1.0172
BHVO1 |  | 0.52 | 718.1142 | 615 | 36 | 11.4 | 23026.3618 | 19720 | 138 | 2.71 | 10388.7191 | 8897 | 96 | 0.168 | 2843.2653 | 2435 | 63 | 11.1176 | 284985.0901 | 244064 | 473 | 45 | 3883.6551 | 3326 | 103 | 121 | 5192.6081 | 4447 | 75 | 136 | 1094.1025 | 937 | 7.2 | 105 | 1442.0668 | 1235 | 7.3 | 11 | 882.755 | 756 | 49 | 403 | 19308.5152 | 16536 | 140 | 27.6 | 1256.408 | 1076 | 73 | 179 | 10026.7429 | 8587 | 138 | 19 | 826.7071 | 708 | 114 | 24364 | 273 | 28449 | 451 | 0.85641 | 0.99908
BHVO1 |  | 0.52 | 619.7342 | 532 | 35 | 11.4 | 23251.6803 | 19960 | 138 | 2.71 | 10357.249 | 8891 | 96 | 0.168 | 2907.6249 | 2496 | 63 | 11.1176 | 286240.2994 | 245718 | 474 | 45 | 3935.079 | 3378 | 104 | 121 | 5256.0913 | 4512 | 77 | 136 | 982.0224 | 843 | 7.2 | 105 | 1392.072 | 1195 | 7.4 | 11 | 905.1381 | 777 | 48 | 403 | 19456.391 | 16702 | 141 | 27.6 | 1304.7035 | 1120 | 74 | 179 | 10202.3154 | 8758 | 139 | 19 | 825.9239 | 709 | 113 | 24722 | 273 | 28799 | 454 | 0.85843 | 0.99672
BHVO1 |  | 0.52 | 797.0474 | 696 | 35 | 11.4 | 23030.7755 | 20111 | 138 | 2.71 | 10215.0324 | 8920 | 96 | 0.168 | 2921.3619 | 2551 | 64 | 11.1176 | 281537.5168 | 245845 | 474 | 45 | 3798.572 | 3317 | 103 | 121 | 5501.4592 | 4804 | 77 | 136 | 991.7285 | 866 | 7 | 105 | 1428.0432 | 1247 | 7.3 | 11 | 918.4368 | 802 | 49 | 403 | 19212.7353 | 16777 | 141 | 27.6 | 1281.4598 | 1119 | 74 | 179 | 9992.8669 | 8726 | 139 | 19 | 1004.3255 | 877 | 115 | 24507 | 274 | 28065 | 455 | 0.87322 | 0.97984
BHVO1 |  | 0.52 | 639.6134 | 573 | 36 | 11.4 | 22349.634 | 20022 | 139 | 2.71 | 9975.9604 | 8937 | 96 | 0.168 | 2742.6356 | 2457 | 64 | 11.1176 | 274862.9919 | 246237 | 475 | 45 | 3569.7797 | 3198 | 104 | 121 | 5128.07 | 4594 | 77 | 136 | 995.6984 | 892 | 7.2 | 105 | 1434.3862 | 1285 | 7.5 | 11 | 905.2818 | 811 | 49 | 403 | 18825.6207 | 16865 | 141 | 27.6 | 1134.1139 | 1016 | 73 | 179 | 9752.7096 | 8737 | 139 | 19 | 858.3992 | 769 | 115 | 24868 | 274 | 27759 | 454 | 0.89585 | 0.95509
BHVO1 |  | 0.52 | 791.9588 | 698 | 36 | 11.4 | 22694.4986 | 20002 | 138 | 2.71 | 10369.2142 | 9139 | 97 | 0.168 | 2841.0671 | 2504 | 65 | 11.1176 | 278801.269 | 245724 | 475 | 45 | 3650.0451 | 3217 | 104 | 121 | 5130.7131 | 4522 | 76 | 136 | 1039.3041 | 916 | 7.1 | 105 | 1461.3796 | 1288 | 7.5 | 11 | 891.8046 | 786 | 49 | 403 | 19329.241 | 17036 | 142 | 27.6 | 1052.9194 | 928 | 73 | 179 | 9813.2546 | 8649 | 139 | 19 | 917.9007 | 809 | 114 | 24567 | 274 | 27874 | 452 | 0.88136 | 0.9708
BCR1 |  | 1.69 | 2240.1298 | 1902 | 52 | 6.95 | 13738.7562 | 11665 | 109 | 2.24 | 9515.2517 | 8079 | 93 | 0.18 | 3192.9505 | 2711 | 67 | 12.1105 | 335090.2079 | 284511 | 510 | 37 | 3877.2384 | 3292 | 106 | 13 | 4543.8595 | 3858 | 71 | 19 | 189.6219 | 161 | 5.2 | 129.5 | 1637.1085 | 1390 | 7.4 | 47.2 | 2849.04 | 2419 | 65 | 330 | 16979.9956 | 14417 | 132 | 38 | 1854.9971 | 1575 | 78 | 190 | 11240.6935 | 9544 | 140 | 14 | 551.1991 | 468 | 114 | 24064 | 274 | 28342 | 457 | 0.84906 | 1.0077
BCR1 |  | 1.69 | 2248.7287 | 1957 | 53 | 6.95 | 13524.546 | 11770 | 109 | 2.24 | 9439.6046 | 8215 | 94 | 0.18 | 3149.599 | 2741 | 67 | 12.1105 | 330769.949 | 287859 | 513 | 37 | 4124.0098 | 3589 | 108 | 13 | 4538.8239 | 3950 | 72 | 19 | 309.0997 | 269 | 5.4 | 129.5 | 1608.6971 | 1400 | 7.4 | 47.2 | 2733.6359 | 2379 | 66 | 330 | 16562.6854 | 14414 | 134 | 38 | 1854.5979 | 1614 | 81 | 190 | 10913.8605 | 9498 | 141 | 14 | 651.5223 | 567 | 115 | 24606 | 275 | 28274 | 451 | 0.87027 | 0.98317
BCR1 |  | 1.69 | 2259.6783 | 1914 | 52 | 6.95 | 14146.0111 | 11982 | 110 | 2.24 | 9571.1661 | 8107 | 92 | 0.18 | 3335.2096 | 2825 | 67 | 12.1105 | 338513.7383 | 286729 | 512 | 37 | 4064.8236 | 3443 | 107 | 13 | 4775.5479 | 4045 | 73 | 19 | 321.1246 | 272 | 5.5 | 129.5 | 1632.7769 | 1383 | 7.3 | 47.2 | 2918.456 | 2472 | 66 | 330 | 17026.6877 | 14422 | 133 | 38 | 1807.5065 | 1531 | 81 | 190 | 11330.2678 | 9597 | 142 | 14 | 552.5232 | 468 | 114 | 24357 | 277 | 28756 | 461 | 0.84702 | 1.0101
BCR1 |  | 1.69 | 2174.9678 | 1907 | 52 | 6.95 | 13553.9156 | 11884 | 109 | 2.24 | 9381.9009 | 8226 | 93 | 0.18 | 3111.333 | 2728 | 67 | 12.1105 | 327812.2385 | 287424 | 513 | 37 | 3761.4283 | 3298 | 108 | 13 | 4559.7909 | 3998 | 73 | 19 | 282.8485 | 248 | 5.5 | 129.5 | 1568.2122 | 1375 | 7.4 | 47.2 | 2880.9484 | 2526 | 66 | 330 | 16446.2692 | 14420 | 133 | 38 | 1685.6856 | 1478 | 79 | 190 | 10631.9086 | 9322 | 141 | 14 | 582.8047 | 511 | 114 | 24673 | 275 | 28140 | 458 | 0.87679 | 0.97585
BCR1 |  | 1.69 | 2230.123 | 1916 | 52 | 6.95 | 13798.5952 | 11855 | 109 | 2.24 | 9713.1402 | 8345 | 94 | 0.18 | 3248.5769 | 2791 | 68 | 12.1105 | 332772.5309 | 285900 | 511 | 37 | 3972.5521 | 3413 | 108 | 13 | 4612.7231 | 3963 | 73 | 19 | 271.1997 | 233 | 5.2 | 129.5 | 1518.9512 | 1305 | 7.4 | 47.2 | 2844.6872 | 2444 | 66 | 330 | 16948.2365 | 14561 | 133 | 38 | 1712.1665 | 1471 | 80 | 190 | 10881.7432 | 9349 | 141 | 14 | 509.8089 | 438 | 113 | 24209 | 276 | 28178 | 460 | 0.85915 | 0.9959
Method #1: Sensitivity Counts/Unit Concentation
Element |  | K2O |  |  |  | CaO |  |  |  | TiO2 |  |  |  | MnO |  |  |  | FeO |  |  |  | Co ppm |  |  |  | Ni ppm |  |  |  | Cu ppm |  |  |  | Zn ppm |  |  |  | Rb ppm |  |  |  | Sr ppm |  |  |  | Y ppm |  |  |  | Zr ppm |  |  |  | Nb ppm |  |  |  | Rh-Ka-Coh Area |  | Rh-Ka-Inc Area
|  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts Peak | %StD | Counts Peak | %StD
BIR1 |  | 4266.6667 | 115.2 | 20.66% |  | 1742.5227 | 23071 | 0.6363% |  | 3088.5417 | 2965 | 2.03% |  | 14936.8421 | 2554.2 | 2.474% |  | 21966.1909 | 225190.4 | 0.2016% |  | 56.4397 | 2901 | 3.392% |  | 27.8675 | 4626 | 1.652% |  | 8.7778 | 1106 | 0.6691% |  | 13.9915 | 993.4 | 0.6845% |  | 1708.1481 | 461.2 | 12.97% |  | 44.7426 | 4832.2 | 1.767% |  | 46.6 | 745.6 | 9.201% |  | 48.4455 | 1065.8 | 8.838% |  | 63.9 | 127.8 | 85.92% |  | 24608.2 | 1.117% | 29487.2 | 1.55%
BHVO1 |  | 1197.6923 | 622.8 | 5.716% |  | 1751.1404 | 19963 | 0.6923% |  | 3305.0923 | 8956.8 | 1.074% |  | 14813.0952 | 2488.6 | 2.564% |  | 22083.7495 | 245517.6 | 0.1931% |  | 73.0489 | 3287.2 | 3.152% |  | 37.8165 | 4575.8 | 1.67% |  | 6.55 | 890.8 | 0.8015% |  | 11.9048 | 1250 | 0.592% |  | 71.4909 | 786.4 | 6.205% |  | 41.6457 | 16783.2 | 0.8401% |  | 38.1087 | 1051.8 | 6.979% |  | 48.5553 | 8691.4 | 1.597% |  | 40.7579 | 774.4 | 14.75% |  | 24605.6 | 1.112% | 28189.2 | 1.608%
BCR1 |  | 1135.6213 | 1919.2 | 2.72% |  | 1702.3309 | 11831.2 | 0.923% |  | 3658.2143 | 8194.4 | 1.137% |  | 15328.8889 | 2759.2 | 2.435% |  | 23655.9814 | 286484.6 | 0.1786% |  | 92.0811 | 3407 | 3.152% |  | 304.8308 | 3962.8 | 1.827% |  | 12.4526 | 236.6 | 2.265% |  | 10.5838 | 1370.6 | 0.5385% |  | 51.8644 | 2448 | 2.688% |  | 43.7782 | 14446.8 | 0.9206% |  | 40.3632 | 1533.8 | 5.203% |  | 49.8 | 9462 | 1.49% |  | 35.0286 | 490.4 | 23.25% |  | 24381.8 | 1.13% | 28338 | 1.614%
Total Counts |  | 1166.6568 |  |  |  | 1731.998 |  |  |  | 3350.6161 |  |  |  | 15026.2754 |  |  |  | 22568.6406 |  |  |  | 73.8566 |  |  |  | 32.842 |  |  |  | 9.2601 |  |  |  | 12.16 |  |  |  | 61.6777 |  |  |  | 43.3888 |  |  |  | 41.6906 |  |  |  | 48.9336 |  |  |  | 37.8932 |  |  |  | 24531.8667 |  | 28671.4667
/wt% or ppm | StDev | 43.8908 |  |  |  | 26.0512 |  |  |  | 287.5518 |  |  |  | 269.2759 |  |  |  | 943.4975 |  |  |  | 17.8344 |  |  |  | 7.035 |  |  |  | 2.9807 |  |  |  | 1.7182 |  |  |  | 13.878 |  |  |  | 1.5848 |  |  |  | 4.3985 |  |  |  | 0.75234 |  |  |  | 4.0512 |  |  |  | 129.968 |  | 710.3527
| %StDev | 3.762% |  |  |  | 1.504% |  |  |  | 8.582% |  |  |  | 1.792% |  |  |  | 4.181% |  |  |  | 24.15% |  |  |  | 21.42% |  |  |  | 32.19% |  |  |  | 14.13% |  |  |  | 22.5% |  |  |  | 3.652% |  |  |  | 10.55% |  |  |  | 1.537% |  |  |  | 10.69% |  |  |  | 0.5298% |  | 2.478%
|  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept
Method #2: |  | 1.0739 | -0.063436 |  |  | 0.92392 | 0.7521 |  |  | 0.94987 | 0.08301 |  |  | 0.42845 | 0.098838 |  |  | 0.68204 | 3.5394 |  |  | -1.8851 | 126.0265 |  |  | 5.6441 | -654.4303 |  |  | 1.201 | -3.2803 |  |  | 1.9217 | -88.1643 |  |  | 1.4257 | -8.8487 |  |  | 1.0845 | -17.7173 |  |  | 1.11 | -2.2833 |  |  | 1.0044 | 0.16268 |  |  | 0.99713 | -0.41338
Slope Correction |  | 0.017411 | 0.017307 |  |  | 0.01878 | 0.20583 |  |  | 0.044447 | 0.09466 |  |  | 0.087379 | 0.015148 |  |  | 0.066587 | 0.74725 |  |  | 0.32187 | 13.9518 |  |  | 0.53697 | 72.0066 |  |  | 0.13182 | 11.934 |  |  | 0.14379 | 14.3212 |  |  | 0.038273 | 0.93089 |  |  | 0.023159 | 6.9025 |  |  | 0.093976 | 2.6001 |  |  | 0.015948 | 2.3982 |  |  | 0.07217 | 1.007
|  | 0.99659 | 0.04372 |  |  | 0.99466 | 0.20732 |  |  | 0.97232 | 0.13215 |  |  | 0.64905 | 0.0032448 |  |  | 0.88975 | 0.27086 |  |  | 0.72517 | 3.3173 |  |  | 0.89472 | 22.3766 |  |  | 0.86461 | 20.9304 |  |  | 0.93216 | 6.7112 |  |  | 0.99072 | 2.0778 |  |  | 0.99411 | 10.3448 |  |  | 0.91476 | 2.8181 |  |  | 0.99673 | 4.7111 |  |  | 0.93624 | 1.9349
|  | 3804.0271 | 13 |  |  | 2420.398 | 13 |  |  | 456.718 | 13 |  |  | 24.0426 | 13 |  |  | 104.9173 | 13 |  |  | 34.3028 | 13 |  |  | 110.4838 | 13 |  |  | 83.0168 | 13 |  |  | 178.6244 | 13 |  |  | 1387.6094 | 13 |  |  | 2192.8951 | 13 |  |  | 139.5066 | 13 |  |  | 3966.5869 | 13 |  |  | 190.8912 | 13
|  | 7.271 | 0.024848 |  |  | 104.0283 | 0.55874 |  |  | 7.976 | 0.22703 |  |  | 0.00025313 | 0.00013687 |  |  | 7.6973 | 0.95374 |  |  | 377.4777 | 143.0556 |  |  | 55320.7232 | 6509.2768 |  |  | 36368.2548 | 5695.0785 |  |  | 8045.3084 | 585.5249 |  |  | 5990.5398 | 56.1232 |  |  | 234672.1415 | 1391.1919 |  |  | 1107.9546 | 103.2454 |  |  | 88034.8101 | 288.5232 |  |  | 714.6636 | 48.6697
Method #3: Traditional |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept
Rh-tube line |  | 0.00078758 | -0.063436 |  |  | 0.00045642 | 0.7521 |  |  | 0.00024256 | 0.08301 |  |  | 2.4396e-05 | 0.098838 |  |  | 2.5858e-05 | 3.5394 |  |  | -0.021839 | 126.0265 |  |  | 0.14704 | -654.4303 |  |  | 0.11097 | -3.2803 |  |  | 0.13522 | -88.1643 |  |  | 0.019778 | -8.8487 |  |  | 0.021386 | -17.7173 |  |  | 0.02278 | -2.2833 |  |  | 0.017563 | 0.16268 |  |  | 0.022515 | -0.41338 |  |  | 24531.8667 | 274.6 | 28671.4667 | 455.8667 | 0.85603 | 1
Corrected Counts |  | 1.2769e-05 | 0.017307 |  |  | 9.2773e-06 | 0.20583 |  |  | 1.135e-05 | 0.09466 |  |  | 4.9755e-06 | 0.015148 |  |  | 2.5244e-06 | 0.74725 |  |  | 0.0037288 | 13.9518 |  |  | 0.013989 | 72.0066 |  |  | 0.01218 | 11.934 |  |  | 0.010117 | 14.3212 |  |  | 0.00053094 | 0.93089 |  |  | 0.00045669 | 6.9025 |  |  | 0.0019287 | 2.6001 |  |  | 0.00027886 | 2.3982 |  |  | 0.0016296 | 1.007
|  | 0.99659 | 0.04372 |  |  | 0.99466 | 0.20732 |  |  | 0.97232 | 0.13215 |  |  | 0.64905 | 0.0032448 |  |  | 0.88975 | 0.27086 |  |  | 0.72517 | 3.3173 |  |  | 0.89472 | 22.3766 |  |  | 0.86461 | 20.9304 |  |  | 0.93216 | 6.7112 |  |  | 0.99072 | 2.0778 |  |  | 0.99411 | 10.3448 |  |  | 0.91476 | 2.8181 |  |  | 0.99673 | 4.7111 |  |  | 0.93624 | 1.9349
|  | 3804.0271 | 13 |  |  | 2420.398 | 13 |  |  | 456.718 | 13 |  |  | 24.0426 | 13 |  |  | 104.9173 | 13 |  |  | 34.3028 | 13 |  |  | 110.4838 | 13 |  |  | 83.0168 | 13 |  |  | 178.6244 | 13 |  |  | 1387.6094 | 13 |  |  | 2192.8951 | 13 |  |  | 139.5066 | 13 |  |  | 3966.5869 | 13 |  |  | 190.8912 | 13
|  | 7.271 | 0.024848 |  |  | 104.0283 | 0.55874 |  |  | 7.976 | 0.22703 |  |  | 0.00025313 | 0.00013687 |  |  | 7.6973 | 0.95374 |  |  | 377.4777 | 143.0556 |  |  | 55320.7232 | 6509.2768 |  |  | 36368.2548 | 5695.0785 |  |  | 8045.3084 | 585.5249 |  |  | 5990.5398 | 56.1232 |  |  | 234672.1415 | 1391.1919 |  |  | 1107.9546 | 103.2454 |  |  | 88034.8101 | 288.5232 |  |  | 714.6636 | 48.6697
Standards as Unknowns |  | K2O | K2O calc | K2O Slope Corr | TRAD | CaO | CaO calc | CaO Slope Corr | TRAD | TiO2 | TiO2 calc | TiO2 Slope Corr | TRAD | MnO | MnO calc | MnO Slope Corr | TRAD | FeO | FeO calc | FeO Slope Corr | TRAD | Co ppm | Co calc | Co slope | TRAD | Ni ppm | Ni calc | Ni slope | TRAD | Cu ppm | Cu calc | Cu slope | TRAD | Zn ppm | Zn calc | Zn slope | TRAD | Rb ppm | Rb calc | Rb slope | TRAD | Sr ppm | Sr calc | Sr slope | TRAD | Y ppm | Y calcc | Y Slope | TRAD | Zr ppm | Zr calc | Zr slope | TRAD | Nb ppm | Nb calc | Nb slope | TRAD
BIR1 | XRF1 | 0.027 | 0.10114 | 0.045173 | 0.045173 | 13.24 | 13.8407 | 13.5397 | 13.5397 | 0.96 | 0.92053 | 0.95712 | 0.9574 | 0.171 | 0.17992 | 0.17605 | 0.17592 | 10.2517 | 10.4096 | 10.6379 | 10.6393 | 51.4 | 40.8224 | 49.2693 | 49.0702 | 166 | 148.2467 | 183.7596 | 182.2937 | 126 | 123.0601 | 144.6124 | 144.5192 | 71 | 86.9888 | 78.7452 | 79.0034 | 0.27 | 7.1655 | 1.3569 | 1.3671 | 108 | 115.2409 | 107.2302 | 107.2619 | 16 | 19.3892 | 19.1781 | 19.2382 | 22 | 25.1279 | 25.3731 | 25.4022 | 2 | 5.1897 | 4.7144 | 4.7614
BIR1 | XRF1 | 0.027 | 0.094401 | 0.037939 | 0.037939 | 13.24 | 13.6333 | 13.3481 | 13.3481 | 0.96 | 0.87427 | 0.91316 | 0.91346 | 0.171 | 0.17995 | 0.17606 | 0.17594 | 10.2517 | 10.2439 | 10.5245 | 10.5262 | 51.4 | 39.4536 | 51.9626 | 51.6507 | 166 | 145.89 | 170.2221 | 168.9922 | 126 | 127.3804 | 149.8108 | 149.7081 | 71 | 90.3161 | 85.2115 | 85.3975 | 0.27 | 6.2747 | 0.086281 | 0.097154 | 108 | 112.1355 | 103.8618 | 103.8942 | 16 | 19.8743 | 19.7208 | 19.7767 | 22 | 22.9695 | 23.2044 | 23.2342 | 2 | 5.0251 | 4.5492 | 4.5973
BIR1 | XRF1 | 0.027 | 0.088127 | 0.031202 | 0.031202 | 13.24 | 13.7238 | 13.4318 | 13.4318 | 0.96 | 0.93406 | 0.96997 | 0.97025 | 0.171 | 0.16661 | 0.1701 | 0.17022 | 10.2517 | 10.2619 | 10.5368 | 10.5385 | 51.4 | 41.1917 | 48.5427 | 48.374 | 166 | 147.7319 | 180.8028 | 179.3884 | 126 | 125.5738 | 147.637 | 147.5384 | 71 | 80.9154 | 66.942 | 67.3321 | 0.27 | 9.3683 | 4.4992 | 4.5077 | 108 | 117.651 | 109.8445 | 109.8757 | 16 | 19.2851 | 19.0617 | 19.1227 | 22 | 21.8725 | 22.1022 | 22.1323 | 2 | 1.5466 | 1.0575 | 1.1287
BIR1 | XRF1 | 0.027 | 0.12724 | 0.073209 | 0.073209 | 13.24 | 13.6332 | 13.348 | 13.348 | 0.96 | 0.88701 | 0.92527 | 0.92556 | 0.171 | 0.16775 | 0.17061 | 0.17071 | 10.2517 | 10.0999 | 10.426 | 10.428 | 51.4 | 40.4197 | 50.0616 | 49.8293 | 166 | 136.5633 | 116.6469 | 116.3511 | 126 | 122.0995 | 143.4566 | 143.3656 | 71 | 79.9373 | 65.0412 | 65.4525 | 0.27 | 8.1603 | 2.7759 | 2.7853 | 108 | 111.3121 | 102.9686 | 103.0011 | 16 | 15.7551 | 15.1128 | 15.2044 | 22 | 20.7788 | 21.0033 | 21.0337 | 2 | 1.7978 | 1.3097 | 1.3792
BIR1 | XRF1 | 0.027 | 0.095034 | 0.038619 | 0.038619 | 13.24 | 13.4553 | 13.1837 | 13.1837 | 0.96 | 0.92045 | 0.95704 | 0.95732 | 0.171 | 0.17722 | 0.17485 | 0.17477 | 10.2517 | 10.1372 | 10.4515 | 10.4535 | 51.4 | 39.4715 | 51.9273 | 51.6168 | 166 | 143.7403 | 157.8739 | 156.8593 | 126 | 114.2385 | 133.998 | 133.9243 | 71 | 80.7215 | 66.5652 | 66.9595 | 0.27 | 7.3554 | 1.6278 | 1.6378 | 108 | 114.5909 | 106.5252 | 106.557 | 16 | 17.4203 | 16.9756 | 17.0528 | 22 | 20.9532 | 21.1786 | 21.2089 | 2 | 3.7581 | 3.2774 | 3.3339
BHVO1 | XRF1 | 0.52 | 0.52666 | 0.50213 | 0.50213 | 11.4 | 11.3752 | 11.2618 | 11.2618 | 2.71 | 2.6529 | 2.6031 | 2.6029 | 0.168 | 0.1619 | 0.168 | 0.1682 | 11.1176 | 10.8043 | 10.9078 | 10.9085 | 45 | 44.9917 | 41.0659 | 41.2105 | 121 | 135.2809 | 109.2808 | 109.1134 | 136 | 101.0931 | 118.1809 | 118.1361 | 105 | 101.4685 | 106.8854 | 106.8293 | 11 | 12.246 | 8.6041 | 8.6103 | 403 | 380.7604 | 395.242 | 395.2191 | 27.6 | 25.7854 | 26.3332 | 26.3378 | 179 | 175.3208 | 176.2796 | 176.2623 | 19 | 18.6668 | 18.2426 | 18.1998
BHVO1 | XRF1 | 0.52 | 0.45451 | 0.42465 | 0.42465 | 11.4 | 11.4865 | 11.3647 | 11.3647 | 2.71 | 2.6448 | 2.5955 | 2.5953 | 0.168 | 0.16556 | 0.16964 | 0.16977 | 11.1176 | 10.8519 | 10.9404 | 10.9409 | 45 | 45.5874 | 39.8938 | 40.0874 | 121 | 136.9349 | 118.7813 | 118.4483 | 136 | 90.7371 | 105.7202 | 105.6981 | 105 | 97.9507 | 100.0488 | 100.0691 | 11 | 12.5565 | 9.047 | 9.053 | 403 | 383.6765 | 398.4051 | 398.3816 | 27.6 | 26.7765 | 27.442 | 27.4379 | 179 | 178.3908 | 179.3641 | 179.3458 | 19 | 18.6492 | 18.2248 | 18.1822
BHVO1 | XRF1 | 0.52 | 0.58455 | 0.5643 | 0.5643 | 11.4 | 11.3774 | 11.2639 | 11.2639 | 2.71 | 2.6085 | 2.5609 | 2.5608 | 0.168 | 0.16635 | 0.16999 | 0.17011 | 11.1176 | 10.6736 | 10.8184 | 10.8193 | 45 | 44.006 | 43.0053 | 43.0686 | 121 | 143.3273 | 155.5015 | 154.5282 | 136 | 91.6339 | 106.7993 | 106.7753 | 105 | 100.4818 | 104.9677 | 104.9331 | 11 | 12.741 | 9.3102 | 9.316 | 403 | 378.8717 | 393.1933 | 393.1707 | 27.6 | 26.2995 | 26.9083 | 26.9084 | 179 | 174.7285 | 175.6845 | 175.6673 | 19 | 22.6774 | 22.2683 | 22.1989
BHVO1 | XRF1 | 0.52 | 0.46909 | 0.44031 | 0.44031 | 11.4 | 11.0409 | 10.953 | 10.953 | 2.71 | 2.5475 | 2.5029 | 2.5028 | 0.168 | 0.15617 | 0.16544 | 0.16575 | 11.1176 | 10.4206 | 10.6453 | 10.6467 | 45 | 41.3555 | 48.2205 | 48.0653 | 121 | 133.5996 | 99.6224 | 99.6235 | 136 | 92.0007 | 107.2406 | 107.2158 | 105 | 100.9281 | 105.8351 | 105.7908 | 11 | 12.5585 | 9.0498 | 9.0558 | 403 | 371.2378 | 384.9128 | 384.8918 | 27.6 | 23.2755 | 23.5255 | 23.5519 | 179 | 170.5293 | 171.4653 | 171.4494 | 19 | 19.3824 | 18.9609 | 18.9134
BHVO1 | XRF1 | 0.52 | 0.58082 | 0.56029 | 0.56029 | 11.4 | 11.2112 | 11.1104 | 11.1104 | 2.71 | 2.6479 | 2.5984 | 2.5982 | 0.168 | 0.16177 | 0.16794 | 0.16815 | 11.1176 | 10.5699 | 10.7475 | 10.7486 | 45 | 42.2853 | 46.3909 | 46.3124 | 121 | 133.6684 | 100.018 | 100.0121 | 136 | 96.0298 | 112.0886 | 112.0549 | 105 | 102.8274 | 109.5263 | 109.4408 | 11 | 12.3715 | 8.7831 | 8.7893 | 403 | 381.1691 | 395.6854 | 395.6624 | 27.6 | 21.6091 | 21.6615 | 21.7023 | 179 | 171.5879 | 172.529 | 172.5128 | 19 | 20.726 | 20.3095 | 20.253
BCR1 | XRF1 | 1.69 | 1.6429 | 1.7008 | 1.7008 | 6.95 | 6.787 | 7.0228 | 7.0228 | 2.24 | 2.4298 | 2.3912 | 2.391 | 0.18 | 0.18181 | 0.1769 | 0.17674 | 12.1105 | 12.7039 | 12.2071 | 12.2041 | 37 | 44.9174 | 41.2122 | 41.3506 | 13 | 118.3794 | 12.1931 | 13.7186 | 19 | 17.5207 | 17.6237 | 17.7627 | 129.5 | 115.1923 | 133.5565 | 133.2025 | 47.2 | 39.5231 | 47.5141 | 47.499 | 330 | 334.8425 | 345.4343 | 345.4208 | 38 | 38.0702 | 40.0757 | 39.9737 | 190 | 196.5471 | 197.6068 | 197.5829 | 14 | 12.4459 | 11.9981 | 11.9968
BCR1 | XRF1 | 1.69 | 1.6492 | 1.7076 | 1.7076 | 6.95 | 6.6812 | 6.925 | 6.925 | 2.24 | 2.4105 | 2.3728 | 2.3727 | 0.18 | 0.17934 | 0.17579 | 0.17568 | 12.1105 | 12.5401 | 12.0951 | 12.0924 | 37 | 47.7762 | 35.5872 | 35.9613 | 13 | 118.2482 | 11.4395 | 12.9782 | 19 | 28.5602 | 30.9068 | 31.0216 | 129.5 | 113.1932 | 129.6713 | 129.3608 | 47.2 | 37.9222 | 45.2305 | 45.2166 | 330 | 326.6132 | 336.5079 | 336.4961 | 38 | 38.062 | 40.0665 | 39.9646 | 190 | 190.8324 | 191.8648 | 191.8427 | 14 | 14.7112 | 14.272 | 14.2556
BCR1 | XRF1 | 1.69 | 1.6572 | 1.7162 | 1.7162 | 6.95 | 6.9882 | 7.2086 | 7.2086 | 2.24 | 2.4441 | 2.4047 | 2.4046 | 0.18 | 0.18991 | 0.18052 | 0.18021 | 12.1105 | 12.8337 | 12.2959 | 12.2926 | 37 | 47.0905 | 36.9363 | 37.2539 | 13 | 124.4154 | 46.8662 | 47.7871 | 19 | 29.6713 | 32.2437 | 32.356 | 129.5 | 114.8875 | 132.9641 | 132.6168 | 47.2 | 40.4861 | 48.8878 | 48.8719 | 330 | 335.7632 | 346.4331 | 346.4194 | 38 | 37.0956 | 38.9854 | 38.8918 | 190 | 198.1134 | 199.1804 | 199.1561 | 14 | 12.4758 | 12.0281 | 12.0266
BCR1 | XRF1 | 1.69 | 1.5951 | 1.6495 | 1.6495 | 6.95 | 6.6957 | 6.9384 | 6.9384 | 2.24 | 2.3958 | 2.3588 | 2.3587 | 0.18 | 0.17716 | 0.17482 | 0.17474 | 12.1105 | 12.428 | 12.0184 | 12.0159 | 37 | 43.5757 | 43.852 | 43.8798 | 13 | 118.7944 | 14.5773 | 16.0612 | 19 | 26.1347 | 27.9883 | 28.1084 | 129.5 | 110.3445 | 124.1352 | 123.8865 | 47.2 | 39.9658 | 48.1456 | 48.1301 | 330 | 324.3175 | 334.0178 | 334.0064 | 38 | 34.5954 | 36.1886 | 36.1167 | 190 | 185.9023 | 186.9114 | 186.8908 | 14 | 13.1596 | 12.7145 | 12.7084
BCR1 | XRF1 | 1.69 | 1.6356 | 1.693 | 1.693 | 6.95 | 6.8166 | 7.0501 | 7.0501 | 2.24 | 2.4804 | 2.4392 | 2.439 | 0.18 | 0.18498 | 0.17831 | 0.17809 | 12.1105 | 12.616 | 12.147 | 12.1441 | 37 | 46.0216 | 39.0396 | 39.2691 | 13 | 120.1734 | 22.4988 | 23.8446 | 19 | 25.0584 | 26.6932 | 26.8157 | 129.5 | 106.8784 | 117.399 | 117.2255 | 47.2 | 39.4627 | 47.428 | 47.413 | 330 | 334.2162 | 344.755 | 344.7416 | 38 | 35.1389 | 36.7966 | 36.72 | 190 | 190.2708 | 191.3006 | 191.2786 | 14 | 11.5114 | 11.06 | 11.0649
|  | K2O | K2O calc | K2O Slope Corr | TRAD | CaO | CaO calc | CaO Slope Corr | TRAD | TiO2 | TiO2 calc | TiO2 Slope Corr | TRAD | MnO | MnO calc | MnO Slope Corr | TRAD | FeO | FeO calc | FeO Slope Corr | TRAD | Co ppm | Co calc | Co slope |  | Ni ppm | Ni calc | Ni slope | TRAD | Cu ppm | Cu calc | Cu slope | TRAD | Zn ppm | Zn calc | Zn slope |  | Rb ppm | Rb calc | Rb slope | TRAD | Sr ppm | Sr calc | Sr slope | TRAD | Y ppm | Y calcc | Y Slope | TRAD | Zr ppm | Zr calc | Zr slope | TRAD | Nb ppm | Nb calc | Nb slope | TRAD
BIR1 | Average | 0.027 | 0.10119 | 0.045229 | 0.045229 | 13.24 | 13.6573 | 13.3703 | 13.3703 | 0.96 | 0.90727 | 0.94451 | 0.9448 | 0.171 | 0.17429 | 0.17353 | 0.17351 | 10.2517 | 10.2305 | 10.5153 | 10.5171 | 51.4 | 40.2718 | 50.3527 | 50.1082 | 166 | 144.4344 | 161.8611 | 160.7769 | 126 | 122.4705 | 143.903 | 143.8111 | 71 | 83.7758 | 72.501 | 72.829 | 0.27 | 7.6648 | 2.0692 | 2.079 | 108 | 114.1861 | 106.0861 | 106.118 | 16 | 18.3448 | 18.0098 | 18.079 | 22 | 22.3404 | 22.5723 | 22.6023 | 2 | 3.4634 | 2.9817 | 3.0401
BHVO1 | Average | 0.52 | 0.52313 | 0.49834 | 0.49834 | 11.4 | 11.2982 | 11.1907 | 11.1907 | 2.71 | 2.6203 | 2.5722 | 2.572 | 0.168 | 0.16235 | 0.1682 | 0.1684 | 11.1176 | 10.6641 | 10.8119 | 10.8128 | 45 | 43.6452 | 43.7153 | 43.7488 | 121 | 136.5622 | 116.6408 | 116.3451 | 136 | 94.2989 | 110.0059 | 109.976 | 105 | 100.7313 | 105.4527 | 105.4126 | 11 | 12.4947 | 8.9588 | 8.9649 | 403 | 379.1431 | 393.4877 | 393.4651 | 27.6 | 24.7492 | 25.1741 | 25.1877 | 179 | 174.1115 | 175.0645 | 175.0475 | 19 | 20.0204 | 19.6012 | 19.5495
BCR1 | Average | 1.69 | 1.636 | 1.6934 | 1.6934 | 6.95 | 6.7938 | 7.029 | 7.029 | 2.24 | 2.4321 | 2.3933 | 2.3932 | 0.18 | 0.18264 | 0.17727 | 0.17709 | 12.1105 | 12.6243 | 12.1527 | 12.1498 | 37 | 45.8763 | 39.3255 | 39.543 | 13 | 120.0022 | 21.515 | 22.8779 | 19 | 25.3891 | 27.0912 | 27.2129 | 129.5 | 112.0992 | 127.5452 | 127.2584 | 47.2 | 39.472 | 47.4412 | 47.4261 | 330 | 331.1505 | 341.4296 | 341.4169 | 38 | 36.5924 | 38.4226 | 38.3334 | 190 | 192.3332 | 193.3728 | 193.3502 | 14 | 12.8608 | 12.4145 | 12.4104
PDF VERSIONS | [6 empty] | PDF VERSIONS | [5 empty] | PDF VERSIONS |  |  |  |  | PDF VERSIONS |  |  |  | PDF VERSIONS |  |  |  | PDF VERSIONS

## STable2C-50 KV XRF1
USGS Standards | Calibration Data | CaO | counts/rh | Ca-Ka Area | Ca-Ka StDev | FeO | counts/rh | Fe-Ka Area | Fe-Ka StDev | Ni ppm | counts/rh | Ni-Ka Area | Ni-Ka StDev | Rb ppm | counts/rh | Rb-Ka Area | Rb-Ka StDev | Sr ppm | counts/rh | Sr-Ka Area | Sr-Ka StDev | Y ppm | counts/rh | Y -Ka Area | Y -Ka StDev | Zr ppm | counts/rh | Zr-Ka Area | Zr-Ka StDev | Nb ppm | counts/rh | Nb-Ka Area | Nb-Ka StDev | Ba ppm | counts/rh | Ba-Ka Area | Ba-Ka StDev |  | Rh co/inc | Rh avg/co/inc
BIR1 |  | 13.24 | 5554.7998 | 4592 | 109 | 10.2517 | 49867.3925 | 41224 | 202 | 166 | 5572.9448 | 4607 | 80 | 0.27 | 112.4992 | 93 | 39 | 108 | 1682.6495 | 1391 | 49 | 16 | 223.7888 | 185 | 35 | 22 | 561.2864 | 464 | 42 | 2 | 19.3547 | 16 | 34 | 7.7 | 1420.1513 | 1174 | 181 |  | 0.82667 | 1.035
BIR1 |  | 13.24 | 5442.2348 | 4524 | 109 | 10.2517 | 48662.5295 | 40452 | 201 | 166 | 5623.8833 | 4675 | 80 | 0.27 | 99.8465 | 83 | 39 | 108 | 1639.6477 | 1363 | 50 | 16 | 235.7821 | 196 | 36 | 22 | 508.8562 | 423 | 41 | 2 | 98.6435 | 82 | 34 | 7.7 | 846.8907 | 704 | 181 |  | 0.83128 | 1.0293
BIR1 |  | 13.24 | 5502.2957 | 4579 | 108 | 10.2517 | 48893.4065 | 40689 | 201 | 166 | 5435.004 | 4523 | 80 | 0.27 | 146.5997 | 122 | 40 | 108 | 1563.3297 | 1301 | 49 | 16 | 263.1585 | 219 | 36 | 22 | 492.6712 | 410 | 41 | 2 | 25.2344 | 21 | 33 | 7.7 | 1051.4323 | 875 | 181 |  | 0.8322 | 1.0281
BIR1 |  | 13.24 | 5050.5042 | 4250 | 107 | 10.2517 | 48317.2827 | 40659 | 201 | 166 | 5462.863 | 4597 | 80 | 0.27 | 250.7427 | 211 | 40 | 108 | 1650.6236 | 1389 | 48 | 16 | 248.366 | 209 | 34 | 22 | 524.0641 | 441 | 42 | 2 | 84.3731 | 71 | 34 | 7.7 | 909.0908 | 765 | 180 |  | 0.8415 | 1.0168
BIR1 |  | 13.24 | 5048.917 | 4247 | 108 | 10.2517 | 48629.8556 | 40906 | 202 | 166 | 5465.0038 | 4597 | 78 | 0.27 | 187.8335 | 158 | 40 | 108 | 1447.9823 | 1218 | 49 | 16 | 274.6173 | 231 | 36 | 22 | 575.3887 | 484 | 41 | 2 | 84.4062 | 71 | 33 | 7.7 | 1028.329 | 865 | 181 |  | 0.84117 | 1.0172
BHVO1 |  | 11.4 | 4485.0028 | 3841 | 104 | 11.1176 | 51585.1224 | 44178 | 209 | 121 | 5084.0152 | 4354 | 79 | 11 | 161.1378 | 138 | 35 | 403 | 5201.9494 | 4455 | 75 | 27.6 | 421.5272 | 361 | 39 | 179 | 3254.2835 | 2787 | 67 | 19 | 425.0302 | 364 | 40 | 139 | 2629.5825 | 2252 | 180 |  | 0.85641 | 0.99908
BHVO1 |  | 11.4 | 4770.3222 | 4095 | 106 | 11.1176 | 52002.9188 | 44641 | 210 | 121 | 5331.8107 | 4577 | 79 | 11 | 158.4283 | 136 | 35 | 403 | 5279.3895 | 4532 | 76 | 27.6 | 391.4111 | 336 | 39 | 179 | 3371.2607 | 2894 | 69 | 19 | 333.1654 | 286 | 39 | 139 | 2577.9543 | 2213 | 181 |  | 0.85843 | 0.99672
BHVO1 |  | 11.4 | 4589.8935 | 4008 | 105 | 11.1176 | 51325.9573 | 44819 | 210 | 121 | 5449.9259 | 4759 | 79 | 11 | 184.3745 | 161 | 34 | 403 | 5210.5827 | 4550 | 76 | 27.6 | 312.635 | 273 | 38 | 179 | 3284.3849 | 2868 | 68 | 19 | 329.8127 | 288 | 39 | 139 | 2162.1055 | 1888 | 181 |  | 0.87322 | 0.97984
BHVO1 |  | 11.4 | 4508.5492 | 4039 | 105 | 11.1176 | 49714.5964 | 44537 | 209 | 121 | 5234.1142 | 4689 | 79 | 11 | 136.183 | 122 | 34 | 403 | 5075.6061 | 4547 | 76 | 27.6 | 360.55 | 323 | 39 | 179 | 3000.4903 | 2688 | 68 | 19 | 334.8761 | 300 | 40 | 139 | 2005.9081 | 1797 | 181 |  | 0.89585 | 0.95509
BHVO1 |  | 11.4 | 4379.6003 | 3860 | 105 | 11.1176 | 51164.1697 | 45094 | 211 | 121 | 5405.289 | 4764 | 80 | 11 | 158.8456 | 140 | 34 | 403 | 5156.8091 | 4545 | 76 | 27.6 | 442.4985 | 390 | 39 | 179 | 3154.2199 | 2780 | 68 | 19 | 342.6527 | 302 | 39 | 139 | 2640.2409 | 2327 | 180 |  | 0.88136 | 0.9708
BCR1 |  | 6.95 | 2929.1287 | 2487 | 98 | 12.1105 | 61233.7484 | 51991 | 225 | 13 | 5368.3027 | 4558 | 80 | 47.2 | 650.1323 | 552 | 40 | 330 | 4563.8817 | 3875 | 71 | 38 | 633.6435 | 538 | 42 | 190 | 3546.2833 | 3011 | 69 | 14 | 232.0219 | 197 | 38 | 681 | 9346.8298 | 7936 | 206 |  | 0.84906 | 1.0077
BCR1 |  | 6.95 | 3115.127 | 2711 | 100 | 12.1105 | 59963.0341 | 52184 | 226 | 13 | 5173.1101 | 4502 | 79 | 47.2 | 621.6465 | 541 | 41 | 330 | 4431.9604 | 3857 | 71 | 38 | 560.7458 | 488 | 42 | 190 | 3477.0838 | 3026 | 69 | 14 | 252.7953 | 220 | 40 | 681 | 8954.6973 | 7793 | 205 |  | 0.87027 | 0.98317
BCR1 |  | 6.95 | 3009.3626 | 2549 | 99 | 12.1105 | 61862.53 | 52399 | 226 | 13 | 5468.5631 | 4632 | 79 | 47.2 | 615.0953 | 521 | 41 | 330 | 4548.8717 | 3853 | 71 | 38 | 545.4396 | 462 | 43 | 190 | 3761.4081 | 3186 | 70 | 14 | 266.8168 | 226 | 39 | 681 | 9486.1625 | 8035 | 205 |  | 0.84702 | 1.0101
BCR1 |  | 6.95 | 2925.4286 | 2565 | 99 | 12.1105 | 60001.5102 | 52609 | 227 | 13 | 5314.8138 | 4660 | 80 | 47.2 | 622.7228 | 546 | 41 | 330 | 4294.0502 | 3765 | 71 | 38 | 541.746 | 475 | 42 | 190 | 3606.3178 | 3162 | 70 | 14 | 297.6752 | 261 | 39 | 681 | 9083.0852 | 7964 | 206 |  | 0.87679 | 0.97585
BCR1 |  | 6.95 | 2990.1806 | 2569 | 99 | 12.1105 | 61347.006 | 52706 | 227 | 13 | 5315.7473 | 4567 | 80 | 47.2 | 608.7444 | 523 | 40 | 330 | 4540.5584 | 3901 | 72 | 38 | 531.9239 | 457 | 42 | 190 | 3412.6935 | 2932 | 69 | 14 | 180.4118 | 155 | 39 | 681 | 9955.2412 | 8553 | 206 |  | 0.85915 | 0.9959
Method #1: Sensitivity Counts/Unit Concentation
Element |  | CaO |  |  |  | FeO |  |  |  | Ni ppm |  |  |  | Rb ppm |  |  |  | Sr ppm |  |  |  | Y ppm |  |  |  | Zr ppm |  |  |  | Nb ppm |  |  |  | Ba ppm
|  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD
BIR1 |  | 335.2266 | 4438.4 | 2.438% |  | 3978.4692 | 40786 | 0.4938% |  | 27.7096 | 4599.8 | 1.731% |  | 494.0741 | 133.4 | 29.69% |  | 12.337 | 1332.4 | 3.678% |  | 13 | 208 | 17.02% |  | 20.2 | 444.4 | 9.316% |  | 26.1 | 52.2 | 64.37% |  | 113.8442 | 876.6 | 20.63%
BHVO1 |  | 348.1228 | 3968.6 | 2.646% |  | 4016.5077 | 44653.8 | 0.4698% |  | 38.2529 | 4628.6 | 1.711% |  | 12.6727 | 139.4 | 24.68% |  | 11.2303 | 4525.8 | 1.675% |  | 12.1957 | 336.6 | 11.53% |  | 15.6615 | 2803.4 | 2.426% |  | 16.2105 | 308 | 12.79% |  | 15.0748 | 2095.4 | 8.619%
BCR1 |  | 370.6763 | 2576.2 | 3.843% |  | 4325.0083 | 52377.8 | 0.4319% |  | 352.6 | 4583.8 | 1.737% |  | 11.3686 | 536.6 | 7.566% |  | 11.6673 | 3850.2 | 1.849% |  | 12.7368 | 484 | 8.719% |  | 16.1232 | 3063.4 | 2.265% |  | 15.1286 | 211.8 | 18.41% |  | 11.83 | 8056.2 | 2.552%
Total Counts | Avg | 351.3419 |  |  |  | 4106.6617 |  |  |  | 32.9813 |  |  |  | 12.0207 |  |  |  | 11.7449 |  |  |  | 12.6442 |  |  |  | 17.3282 |  |  |  | 15.6695 |  |  |  | 13.4524
/wt% or ppm | StDev | 17.9427 |  |  |  | 190.0477 |  |  |  | 7.4552 |  |  |  | 0.92213 |  |  |  | 0.55745 |  |  |  | 0.4101 |  |  |  | 2.4977 |  |  |  | 0.76506 |  |  |  | 2.2945
| %StDev | 5.107% |  |  |  | 4.628% |  |  |  | 22.6% |  |  |  | 7.671% |  |  |  | 4.746% |  |  |  | 3.243% |  |  |  | 14.41% |  |  |  | 4.882% |  |  |  | 17.06%
Method #2: |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept
Slope Correction |  | 1.0968 | -0.91978 |  |  | 0.65816 | 3.8047 |  |  | 8.4777 | -1083.3391 |  |  | 1.2339 | -8.1184 |  |  | 1.105 | -21.742 |  |  | 0.95564 | 1.3854 |  |  | 1.1408 | -6.8503 |  |  | 0.99395 | -0.29981 |  |  | 1.2511 | -64.1018
|  | 0.041814 | 0.44771 |  |  | 0.055789 | 0.62644 |  |  | 4.1624 | 581.1687 |  |  | 0.08208 | 2.2446 |  |  | 0.020076 | 5.9398 |  |  | 0.066545 | 1.8938 |  |  | 0.023796 | 3.2752 |  |  | 0.072803 | 1.0066 |  |  | 0.02415 | 8.6333
|  | 0.97727 | 0.42222 |  |  | 0.89689 | 0.25865 |  |  | 0.20589 | 60.6835 |  |  | 0.93389 | 5.4757 |  |  | 0.99475 | 9.6395 |  |  | 0.92959 | 2.529 |  |  | 0.9931 | 6.7607 |  |  | 0.92299 | 2.0997 |  |  | 0.99407 | 23.796
|  | 688.0223 | 16 |  |  | 139.1749 | 16 |  |  | 4.1484 | 16 |  |  | 226.0052 | 16 |  |  | 3032.6295 | 16 |  |  | 211.2538 | 16 |  |  | 2302.8415 | 16 |  |  | 191.7685 | 16 |  |  | 2683.6472 | 16
|  | 122.6521 | 2.8523 |  |  | 9.3108 | 1.0704 |  |  | 15276.2861 | 58919.7139 |  |  | 6776.2707 | 479.7249 |  |  | 281789.2939 | 1486.7061 |  |  | 1351.1093 | 102.3307 |  |  | 105256.6831 | 731.3169 |  |  | 845.46 | 70.54 |  |  | 1519613.1722 | 9059.9878
Method #3: Traditional |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept
Rh-tube line Corrected Counts |  | 0.0026596 | -0.87115 |  |  | 0.00013672 | 3.8268 |  |  | 0.19957 | -973.783 |  |  | 0.087487 | -8.009 |  |  | 0.080499 | -21.742 |  |  | 0.064668 | 1.3854 |  |  | 0.056329 | -6.8503 |  |  | 0.054274 | -0.29981 |  |  | 0.079494 | -63.793
|  | 0.00012237 | 0.53808 |  |  | 1.304e-05 | 0.7028 |  |  | 0.11555 | 621.9189 |  |  | 0.0066419 | 2.5534 |  |  | 0.001699 | 6.9005 |  |  | 0.0053386 | 2.2466 |  |  | 0.0013883 | 3.8704 |  |  | 0.0047303 | 1.1998 |  |  | 0.0018361 | 10.3225
|  | 0.97322 | 0.46419 |  |  | 0.89424 | 0.26529 |  |  | 0.18664 | 62.1971 |  |  | 0.93029 | 5.694 |  |  | 0.99424 | 10.2249 |  |  | 0.91861 | 2.7537 |  |  | 0.99217 | 7.2958 |  |  | 0.91012 | 2.2972 |  |  | 0.99311 | 25.9793
|  | 472.3926 | 13 |  |  | 109.9165 | 13 |  |  | 2.983 | 13 |  |  | 173.4997 | 13 |  |  | 2244.9358 | 13 |  |  | 146.7293 | 13 |  |  | 1646.3198 | 13 |  |  | 131.645 | 13 |  |  | 1874.4701 | 13
|  | 101.7859 | 2.8011 |  |  | 7.736 | 0.91495 |  |  | 11539.7353 | 50290.2647 |  |  | 5625.1791 | 421.4839 |  |  | 234704.2055 | 1359.1278 |  |  | 1112.6232 | 98.5768 |  |  | 87631.361 | 691.9723 |  |  | 694.7286 | 68.6047 |  |  | 1265120.3193 | 8773.9807
Standards as Unknowns |  | CaO | CaO calc | CaO slope | TRAD | FeO | FeO calc | FeO slope | TRAD | Ni ppm | Ni calc | Ni slope | TRAD | Rb ppm | Rb calc | Rb slope | TRAD | Sr ppm | Sr calc | Sr slope | TRAD | Y ppm | Y calc | Y slope | TRAD | Zr ppm | Zr calc | Zr slope | TRAD | Nb ppm | Nb calc | Nb slope | TRAD | Ba ppm | Ba calc | Ba slope | TRAD
BIR1 | XRF1 | 13.24 | 13.5276 | 13.9171 | 13.9025 | 10.2517 | 10.3898 | 10.6429 | 10.6444 | 166 | 144.5766 | 142.344 | 138.4352 | 0.27 | 8.0076 | 1.7625 | 1.8332 | 108 | 122.5819 | 113.7099 | 113.7099 | 16 | 15.1436 | 15.8573 | 15.8573 | 22 | 27.7148 | 24.7667 | 24.7667 | 2 | 1.0568 | 0.75063 | 0.75063 | 7.7 | 90.3267 | 48.9022 | 49.1008
BIR1 | XRF1 | 13.24 | 13.2534 | 13.6164 | 13.6031 | 10.2517 | 10.1388 | 10.4777 | 10.4797 | 166 | 145.8981 | 153.5471 | 148.6012 | 0.27 | 7.107 | 0.65117 | 0.72626 | 108 | 119.4492 | 110.2483 | 110.2483 | 16 | 15.9552 | 16.6329 | 16.6329 | 22 | 25.1259 | 21.8133 | 21.8133 | 2 | 5.3863 | 5.0539 | 5.0539 | 7.7 | 53.8652 | 3.2868 | 3.5299
BIR1 | XRF1 | 13.24 | 13.3997 | 13.7769 | 13.7629 | 10.2517 | 10.1869 | 10.5093 | 10.5112 | 166 | 140.9981 | 112.006 | 110.9057 | 0.27 | 10.4348 | 4.7575 | 4.8165 | 108 | 113.8894 | 104.1048 | 104.1048 | 16 | 17.8077 | 18.4032 | 18.4032 | 22 | 24.3268 | 20.9016 | 20.9016 | 2 | 1.3779 | 1.0697 | 1.0697 | 7.7 | 66.8748 | 19.5625 | 19.7898
BIR1 | XRF1 | 13.24 | 12.2994 | 12.5701 | 12.5613 | 10.2517 | 10.0669 | 10.4303 | 10.4325 | 166 | 141.7208 | 118.1332 | 116.4656 | 0.27 | 17.8476 | 13.9045 | 13.9277 | 108 | 120.2488 | 111.1318 | 111.1318 | 16 | 16.8067 | 17.4466 | 17.4466 | 22 | 25.8769 | 22.6699 | 22.6699 | 2 | 4.6071 | 4.2794 | 4.2794 | 7.7 | 57.8214 | 8.2362 | 8.4744
BIR1 | XRF1 | 13.24 | 12.2956 | 12.5659 | 12.5571 | 10.2517 | 10.132 | 10.4732 | 10.4752 | 166 | 141.7764 | 118.604 | 116.8929 | 0.27 | 13.3698 | 8.3791 | 8.4239 | 108 | 105.4863 | 94.8194 | 94.8194 | 16 | 18.5831 | 19.1443 | 19.1443 | 22 | 28.4111 | 25.561 | 25.561 | 2 | 4.6089 | 4.2812 | 4.2812 | 7.7 | 65.4054 | 17.7242 | 17.9532
BHVO1 | XRF1 | 11.4 | 10.9223 | 11.0597 | 11.0573 | 11.1176 | 10.7477 | 10.8784 | 10.8792 | 121 | 131.8925 | 34.8115 | 40.8572 | 11 | 11.4696 | 6.0344 | 6.0884 | 403 | 378.9649 | 397.0108 | 397.0108 | 27.6 | 28.5244 | 28.6446 | 28.6446 | 179 | 160.6877 | 176.4618 | 176.4618 | 19 | 23.2083 | 22.7681 | 22.7681 | 139 | 167.2508 | 145.1388 | 145.2435
BHVO1 | XRF1 | 11.4 | 11.6171 | 11.8218 | 11.8161 | 11.1176 | 10.8348 | 10.9357 | 10.9364 | 121 | 138.321 | 89.3103 | 90.3109 | 11 | 11.2768 | 5.7965 | 5.8514 | 403 | 384.6064 | 403.2446 | 403.2446 | 27.6 | 26.4864 | 26.697 | 26.697 | 179 | 166.4637 | 183.051 | 183.051 | 19 | 18.1922 | 17.7822 | 17.7822 | 139 | 163.967 | 141.0306 | 141.1394
BHVO1 | XRF1 | 11.4 | 11.1777 | 11.3398 | 11.3362 | 11.1176 | 10.6937 | 10.8429 | 10.8438 | 121 | 141.3852 | 115.2879 | 113.8837 | 11 | 13.1236 | 8.0753 | 8.1213 | 403 | 379.5938 | 397.7057 | 397.7057 | 27.6 | 21.1557 | 21.6028 | 21.6028 | 179 | 162.174 | 178.1574 | 178.1574 | 19 | 18.0091 | 17.6003 | 17.6003 | 139 | 137.5176 | 107.9408 | 108.0818
BHVO1 | XRF1 | 11.4 | 10.9796 | 11.1226 | 11.1199 | 11.1176 | 10.358 | 10.6219 | 10.6235 | 121 | 135.7865 | 67.8234 | 70.8132 | 11 | 9.6934 | 3.8426 | 3.9052 | 403 | 369.7607 | 386.8402 | 386.8402 | 27.6 | 24.3981 | 24.7013 | 24.7013 | 179 | 148.156 | 162.1657 | 162.1657 | 19 | 18.2856 | 17.8751 | 17.8751 | 139 | 127.5829 | 95.5119 | 95.6651
BHVO1 | XRF1 | 11.4 | 10.6656 | 10.7781 | 10.7769 | 11.1176 | 10.66 | 10.8207 | 10.8217 | 121 | 140.2272 | 105.4707 | 104.9753 | 11 | 11.3065 | 5.8331 | 5.8879 | 403 | 375.6764 | 393.377 | 393.377 | 27.6 | 29.9435 | 30.0007 | 30.0007 | 179 | 155.7468 | 170.8252 | 170.8252 | 19 | 18.7102 | 18.2972 | 18.2972 | 139 | 167.9287 | 145.9869 | 146.0908
BCR1 | XRF1 | 6.95 | 7.1333 | 6.9039 | 6.9192 | 12.1105 | 12.758 | 12.2015 | 12.1983 | 13 | 139.2677 | 97.3361 | 97.5938 | 47.2 | 46.2757 | 48.9833 | 48.8689 | 330 | 332.4813 | 345.6468 | 345.6468 | 38 | 42.8781 | 42.3616 | 42.3616 | 190 | 175.1058 | 192.91 | 192.91 | 14 | 12.6693 | 12.2928 | 12.2928 | 681 | 594.4915 | 679.6422 | 679.2257
BCR1 | XRF1 | 6.95 | 7.5862 | 7.4007 | 7.4139 | 12.1105 | 12.4932 | 12.0273 | 12.0246 | 13 | 134.2039 | 54.4065 | 58.6383 | 47.2 | 44.2481 | 46.4813 | 46.3768 | 330 | 322.8707 | 335.0273 | 335.0273 | 38 | 37.9452 | 37.6475 | 37.6475 | 190 | 171.6889 | 189.012 | 189.012 | 14 | 13.8036 | 13.4203 | 13.4203 | 681 | 569.5505 | 648.4395 | 648.0534
BCR1 | XRF1 | 6.95 | 7.3287 | 7.1182 | 7.1326 | 12.1105 | 12.889 | 12.2878 | 12.2843 | 13 | 141.8687 | 119.3869 | 117.6032 | 47.2 | 43.7818 | 45.9059 | 45.8036 | 330 | 331.3878 | 344.4385 | 344.4385 | 38 | 36.9094 | 36.6577 | 36.6577 | 190 | 185.7281 | 205.0278 | 205.0278 | 14 | 14.5693 | 14.1813 | 14.1813 | 681 | 603.3535 | 690.7292 | 690.3018
BCR1 | XRF1 | 6.95 | 7.1243 | 6.894 | 6.9094 | 12.1105 | 12.5013 | 12.0326 | 12.0299 | 13 | 137.88 | 85.5721 | 86.9188 | 47.2 | 44.3247 | 46.5759 | 46.4709 | 330 | 312.8239 | 323.9256 | 323.9256 | 38 | 36.6595 | 36.4188 | 36.4188 | 190 | 178.0702 | 196.2917 | 196.2917 | 14 | 16.2542 | 15.8561 | 15.8561 | 681 | 577.7164 | 658.6556 | 658.2595
BCR1 | XRF1 | 6.95 | 7.282 | 7.067 | 7.0816 | 12.1105 | 12.7816 | 12.2171 | 12.2138 | 13 | 137.9043 | 85.7774 | 87.1051 | 47.2 | 43.3298 | 45.3481 | 45.248 | 330 | 330.7822 | 343.7693 | 343.7693 | 38 | 35.9948 | 35.7837 | 35.7837 | 190 | 168.5095 | 185.3849 | 185.3849 | 14 | 9.8512 | 9.4918 | 9.4918 | 681 | 633.1886 | 728.0546 | 727.5908
Standards as Unknowns | Average Compositions | CaO | CaO calc | CaO slope | TRAD | FeO | FeO calc | FeO slope | TRAD | Ni ppm | Ni calc | Ni slope | TRAD | Rb ppm | Rb calc | Rb slope | TRAD | Sr ppm | Sr calc | Sr slope | TRAD | Y ppm | Y calc | Y slope | TRAD | Zr ppm | Zr calc | Zr slope | TRAD | Nb ppm | Nb calc | Nb slope | TRAD | Ba ppm | Ba calc | Ba slope | TRAD
BIR-1 |  | 13.24 | 12.9551 | 13.2893 | 13.2774 | 10.2517 | 10.1829 | 10.5067 | 10.5086 | 166 | 142.994 | 128.9269 | 126.2601 | 0.27 | 11.3533 | 5.891 | 5.9455 | 108 | 116.3311 | 106.8028 | 106.8028 | 16 | 16.8593 | 17.4969 | 17.4969 | 22 | 26.2911 | 23.1425 | 23.1425 | 2 | 3.4074 | 3.087 | 3.087 | 7.7 | 66.8587 | 19.5424 | 19.7696
BHVO-1 |  | 11.4 | 11.0725 | 11.2244 | 11.2213 | 11.1176 | 10.6588 | 10.8199 | 10.8209 | 121 | 137.5225 | 82.5408 | 84.1681 | 11 | 11.374 | 5.9164 | 5.9708 | 403 | 377.7204 | 395.6357 | 395.6357 | 27.6 | 26.1016 | 26.3293 | 26.3293 | 179 | 158.6456 | 174.1322 | 174.1322 | 19 | 19.2811 | 18.8646 | 18.8646 | 139 | 152.8494 | 127.1218 | 127.2441
BCR-1 |  | 6.95 | 7.2909 | 7.0768 | 7.0913 | 12.1105 | 12.6846 | 12.1532 | 12.1502 | 13 | 138.2249 | 88.4958 | 89.5718 | 47.2 | 44.392 | 46.6589 | 46.5536 | 330 | 326.0692 | 338.5615 | 338.5615 | 38 | 38.0774 | 37.7739 | 37.7739 | 190 | 175.8205 | 193.7253 | 193.7253 | 14 | 13.4295 | 13.0484 | 13.0484 | 681 | 595.6601 | 681.1042 | 680.6862
PDF VERSIONS | [5 empty] | PDF VERSIONS | [5 empty] | PDF VERSIONS |  |  |  |  | PDF VERSIONS |  |  |  |  | PDF VERSIONS |  |  |  |  | PDF VERSIONS | [5 empty] | PDF VERSIONS |  |  |  |  | PDF VERSIONS

## All XRF1 USGS
10 KV | Element | MgO | Al2O3 | SiO2 | P2O5 | K2O | CaO | TiO2 | Cr ppm | MnO | FeO | [10 empty] | Rh-La Coh
Total Counts | Avg | 277.488 | 1756.2618 | 3838.1937 | 6331.0256 | 24984.497 | 31727.8843 | 31259.8049 | 11.1585 | 35547.7527 | 34621.4545 | [10 empty] | 31707.0667
/wt% or ppm | stdev | 28.2135 | 37.1423 | 152.852 | 446.3838 | 1873.1217 | 526.3352 | 1718.3855 | 0.73695 | 1915.2494 | 628.9348 | [10 empty] | 309.7987
| %std | 10.17% | 2.115% | 3.982% | 7.051% | 7.497% | 1.659% | 5.497% | 6.604% | 5.388% | 1.817% | [10 empty] | 0.9771%
30 KV | Element |  |  |  |  | K2O | CaO | TiO2 |  | MnO | FeO | Co ppm | Ni ppm | Cu ppm | Zn ppm | Rb ppm | Sr ppm | Y ppm | Zr ppm | Nb ppm |  | Rh-Ka-Coh | Rh-Ka-Inc
Total Counts | Avg |  |  |  |  | 1166.6568 | 1731.998 | 3350.6161 |  | 15026.2754 | 22568.6406 | 73.8566 | 32.842 | 9.2601 | 12.16 | 61.6777 | 43.3888 | 41.6906 | 48.9336 | 37.8932 |  | 24531.8667 | 28671.4667
/wt% or ppm | StDev |  |  |  |  | 43.8908 | 26.0512 | 287.5518 |  | 269.2759 | 943.4975 | 17.8344 | 7.035 | 2.9807 | 1.7182 | 13.878 | 1.5848 | 4.3985 | 0.75234 | 4.0512 |  | 129.968 | 710.3527
| %StDev |  |  |  |  | 3.762% | 1.504% | 8.582% |  | 1.792% | 4.181% | 24.15% | 21.42% | 32.19% | 14.13% | 22.5% | 3.652% | 10.55% | 1.537% | 10.69% |  | 0.5298% | 2.478%
50 V | Element | [5 empty] | CaO |  |  |  | FeO |  | Ni ppm |  |  | Rb ppm | Sr ppm | Y ppm | Zr ppm | Nb ppm | Ba ppm
Total Counts | Avg | [5 empty] | 351.3419 |  |  |  | 4106.6617 |  | 32.9813 |  |  | 12.0207 | 11.7449 | 12.6442 | 17.3282 | 15.6695 | 13.4524
/wt% or ppm | StDev | [5 empty] | 17.9427 |  |  |  | 190.0477 |  | 7.4552 |  |  | 0.92213 | 0.55745 | 0.4101 | 2.4977 | 0.76506 | 2.2945
| %StDev | [5 empty] | 5.107% |  |  |  | 4.628% |  | 22.6% |  |  | 7.671% | 4.746% | 3.243% | 14.41% | 4.882% | 17.06%

## AVG USGS
Table ##-A. USGS standards treated as unknowns using average counts on each standard. Columns compare USGS Reference Values to Calculated Values (mass absorption corrected), slope-corrected concentrations, and concentrations derived by traditional regression of tube-line normalized count data.
|  | BIR1 |  |  |  | BHVO1 |  |  |  | BCR1
10 KV XRF-1 | Reference value | Calculated | Slope Corrected | Traditional Regressed | Reference value | Calculated | Slope Corrected | Traditional Regressed | Reference value | Calculated | Slope Corrected | Traditional Regressed
SiO2 | 47.96 | 45.9336 | 47.8338 | 47.8384 | 49.79 | 49.468 | 49.8862 | 49.8874 | 54.11 | 56.7241 | 54.0999 | 54.0942
TiO2 | 0.96 | 0.90732 | 0.94342 | 0.94345 | 2.74 | 2.6916 | 2.631 | 2.6309 | 2.24 | 2.3794 | 2.3356 | 2.3356
Al2O3 | 15.5 | 15.66 | 15.2736 | 15.2709 | 13.69 | 13.3387 | 13.5815 | 13.5834 | 13.65 | 13.8371 | 13.9448 | 13.9457
FeO | 10.17 | 10.2123 | 10.4103 | 10.4107 | 11.09 | 10.841 | 10.901 | 10.9011 | 12.1105 | 12.4652 | 12.1685 | 12.1679
MnO | 0.171 | 0.16707 | 0.17093 | 0.17096 | 0.168 | 0.16011 | 0.16862 | 0.1687 | 0.18 | 0.19281 | 0.17945 | 0.17934
MgO | 9.68 | 10.3972 | 9.7405 | 9.7346 | 7.21 | 6.6389 | 7.068 | 7.0675 | 3.48 | 1.7358 | 3.5814 | 3.5879
CaO | 13.3 | 13.4402 | 13.3136 | 13.3134 | 11.43 | 11.3429 | 11.2906 | 11.2905 | 6.95 | 6.88 | 6.9859 | 6.986
K2O | 0.03 | 0.0026847 | 0.046607 | 0.046801 | 0.526 | 0.48728 | 0.49286 | 0.49299 | 1.69 | 1.7951 | 1.6973 | 1.6972
P2O5 | 0.021 | 0.060441 | 0.055934 | 0.056137 | 0.277 | 0.25669 | 0.25 | 0.24998 | 0.36 | 0.38112 | 0.37305 | 0.37289
Na2O | 1.82 | 1.82 | 1.82 | 1.82 | 2.31 | 2.31 | 2.31 | 2.31 | 3.27 | 3.27 | 3.27 | 3.27
Cr ppm | 382 | 400.6581 | 378.1295 | 401.7995 | 289 | 272.6223 | 292.8753 | 248.7917 | 16 | 94.902 | 174.538 | 36.4088
| 99.612 | 96.7809 | 99.6088 | 99.6054 | 99.231 | 97.5352 | 98.5797 | 98.5825 | 98.0405 | 99.6608 | 98.6359 | 98.6369
30 KV XRF-1 | Reference value | Calculated | Slope Corrected | Traditional Regressed | Reference value | Calculated | Slope Corrected | Traditional Regressed | Reference value | Calculated | Slope Corrected | Traditional Regressed
K2O | 0.027 | 0.10119 | 0.044874 | 0.045229 | 0.52 | 0.52313 | 0.49821 | 0.49834 | 1.69 | 1.636 | 1.6939 | 1.6934
CaO | 13.24 | 13.6573 | 13.3713 | 13.3703 | 11.4 | 11.2982 | 11.191 | 11.1907 | 6.95 | 6.7938 | 7.0278 | 7.029
TiO2 | 0.96 | 0.90727 | 0.94451 | 0.9448 | 2.71 | 2.6203 | 2.5722 | 2.572 | 2.24 | 2.4321 | 2.3933 | 2.3932
MnO | 0.171 | 0.17429 | 0.17353 | 0.17351 | 0.168 | 0.16235 | 0.1682 | 0.1684 | 0.18 | 0.18264 | 0.17727 | 0.17709
FeO | 10.2517 | 10.2305 | 10.5153 | 10.5171 | 11.1176 | 10.6641 | 10.8119 | 10.8128 | 12.1105 | 12.6243 | 12.1527 | 12.1498
Co ppm | 51.4 | 40.2718 | 50.3527 | 50.1082 | 45 | 43.6452 | 43.7153 | 43.7488 | 37 | 45.8763 | 39.3255 | 39.543
Ni ppm | 166 | 144.4344 | 161.8611 | 160.7769 | 121 | 136.5622 | 116.6408 | 116.3451 | 13 | 120.0022 | 21.515 | 22.8779
Cu ppm | 126 | 122.4705 | 143.903 | 143.8111 | 136 | 94.2989 | 110.0059 | 109.976 | 19 | 25.3891 | 27.0912 | 27.2129
Zn ppm | 71 | 83.7758 | 72.501 | 72.829 | 105 | 100.7313 | 105.4527 | 105.4126 | 129.5 | 112.0992 | 127.5452 | 127.2584
Rb ppm | 0.27 | 7.6648 | 2.0692 | 2.079 | 11 | 12.4947 | 8.9588 | 8.9649 | 47.2 | 39.472 | 47.4412 | 47.4261
Sr ppm | 108 | 114.1861 | 106.0861 | 106.118 | 403 | 379.1431 | 393.4877 | 393.4651 | 330 | 331.1505 | 341.4296 | 341.4169
Y ppm | 16 | 18.3448 | 18.0098 | 18.079 | 27.6 | 24.7492 | 25.1741 | 25.1877 | 38 | 36.5924 | 38.4226 | 38.3334
Zr ppm | 22 | 22.3404 | 22.5723 | 22.6023 | 179 | 174.1115 | 175.0645 | 175.0475 | 190 | 192.3332 | 193.3728 | 193.3502
Nb ppm | 2 | 3.4634 | 2.9817 | 3.0401 | 19 | 20.0204 | 19.6012 | 19.5495 | 14 | 12.8608 | 12.4145 | 12.4104
50 KV XRF-1 | Reference value | Calculated | Slope Corrected | Traditional Regressed | Reference value | Calculated | Slope Corrected | Traditional Regressed | Reference value | Calculated | Slope Corrected | Traditional Regressed
CaO | 13.24 | 12.9551 | 13.2893 | 13.2774 | 11.4 | 11.0725 | 11.2244 | 11.2213 | 6.95 | 7.2909 | 7.0768 | 7.0913
FeO | 10.2517 | 10.1829 | 10.5067 | 10.5086 | 11.1176 | 10.6588 | 10.8199 | 10.8209 | 12.1105 | 12.6846 | 12.1532 | 12.1502
Ni ppm | 166 | 142.994 | 128.9269 | 126.2601 | 121 | 137.5225 | 82.5408 | 84.1681 | 13 | 138.2249 | 88.4958 | 89.5718
Rb ppm | 0.27 | 11.3533 | 5.891 | 5.9455 | 11 | 11.374 | 5.9164 | 5.9708 | 47.2 | 44.392 | 46.6589 | 46.5536
Sr ppm | 108 | 116.3311 | 106.7195 | 106.8028 | 403 | 377.7204 | 395.6994 | 395.6357 | 330 | 326.0692 | 338.5962 | 338.5615
Y ppm | 16 | 16.8593 | 17.3801 | 17.4969 | 27.6 | 26.1016 | 26.3194 | 26.3293 | 38 | 38.0774 | 37.9024 | 37.7739
Zr ppm | 22 | 26.2911 | 23.0413 | 23.1425 | 179 | 158.6456 | 174.18 | 174.1322 | 190 | 175.8205 | 193.7924 | 193.7253
Nb ppm | 2 | 3.4074 | 2.9638 | 3.087 | 19 | 19.2811 | 18.9673 | 18.8646 | 14 | 13.4295 | 13.0679 | 13.0484
Ba ppm | 7.7 | 66.8587 | 19.5424 | 19.7696 | 139 | 152.8494 | 127.1218 | 127.2441 | 681 | 595.6601 | 681.1042 | 680.6862