## CAPTION SuppTab3
Supplemental Table 3.
Calibration Data for XRF2. (A) 10 KV, (B) 30 KV, (C) 50 KV. Each includes calibration data for USGS Standards on XRF1, Calibration Method #1: Sensitivity/First Approximation, Slope Corrections; Calibration Method #2: Traditional regression calibration; and standards data treated as unknowns, with calculated (first approximation), slope-corrected, and traditional regression analysis compared to reference values for each individual calibration count.

## STable3A-10 KV XRF2
USGS Standards | Calibration Data | MgO | Counts /Rh | Mg-Ka Area | Mg-Ka AreaStd | Al2O3 | Counts /Rh | Al-Ka Area | Al-Ka AreaStd | SiO2 | Counts /Rh | Si-Ka Area | Si-Ka AreaStd | P2O5 | Counts /Rh | P -Ka Area | P -Ka AreaStd | K2O | Counts /Rh | K -Ka Area | K -Ka AreaStd | CaO | Counts /Rh | Ca-Ka Area | Ca-Ka AreaStd | TiO2 | Counts /Rh | Ti-Ka Area | Ti-Ka AreaStd | Cr ppm | Counts /Rh | Cr-Ka Area | Cr-Ka AreaStd | MnO | Counts /Rh | Mn-Ka Area | Mn-Ka AreaStd | FeO | Counts /Rh | Fe-Ka Area | Fe-Ka AreaStd |  | Rh-La Coh Area | Rh-La-Coh AreaStd | ratio rh | Ka 30kv co/inc
BIR1 |  | 9.68 | 0.086701 | 3100 | 77 | 15.35 | 0.77209 | 27606 | 177 | 47.77 | 5.3889 | 192680 | 440 | 0.046 | 0.0073556 | 263 | 67 | 0.027 | -0.0037477 | -134 | 117 | 13.24 | 12.4739 | 446006 | 651 | 0.96 | 0.85132 | 30439 | 217 | 382 | 0.14362 | 5135 | 193 | 0.171 | 0.19144 | 6845 | 183 | 10.2517 | 10.5354 | 376695 | 617 |  | 35755 | 208 | 1.0085 | 0.9934
BIR1 |  | 9.68 | 0.088233 | 3189 | 78 | 15.35 | 0.77301 | 27939 | 179 | 47.77 | 5.3158 | 192129 | 439 | 0.046 | 0.006585 | 238 | 68 | 0.027 | -0.0030988 | -112 | 117 | 13.24 | 12.3478 | 446288 | 650 | 0.96 | 0.83081 | 30028 | 216 | 382 | 0.13981 | 5053 | 176 | 0.171 | 0.18416 | 6656 | 180 | 10.2517 | 10.442 | 377407 | 615 |  | 36143 | 168 | 0.99769 | 0.97032
BIR1 |  | 9.68 | 0.090349 | 3226 | 78 | 15.35 | 0.782 | 27922 | 178 | 47.77 | 5.4 | 192811 | 440 | 0.046 | 0.0079538 | 284 | 67 | 0.027 | -0.0066375 | -237 | 117 | 13.24 | 12.5324 | 447482 | 652 | 0.96 | 0.84902 | 30315 | 216 | 382 | 0.15028 | 5366 | 195 | 0.171 | 0.18241 | 6513 | 181 | 10.2517 | 10.5537 | 376832 | 617 |  | 35706 | 208 | 1.0099 | 1.0006
BIR1 |  | 9.68 | 0.087871 | 3142 | 77 | 15.35 | 0.78508 | 28072 | 179 | 47.77 | 5.4015 | 193141 | 441 | 0.046 | 0.0071594 | 256 | 70 | 0.027 | -0.0079705 | -285 | 117 | 13.24 | 12.5045 | 447123 | 650 | 0.96 | 0.84705 | 30288 | 217 | 382 | 0.14954 | 5347 | 177 | 0.171 | 0.18438 | 6593 | 181 | 10.2517 | 10.5742 | 378103 | 616 |  | 35757 | 168 | 1.0085 | 0.99036
BIR1 |  | 9.68 | 0.083982 | 3008 | 76 | 15.35 | 0.7521 | 26938 | 176 | 47.77 | 5.2345 | 187484 | 434 | 0.046 | 0.0075662 | 271 | 68 | 0.027 | -0.010945 | -392 | 116 | 13.24 | 12.1819 | 436319 | 643 | 0.96 | 0.82215 | 29447 | 214 | 382 | 0.13692 | 4904 | 174 | 0.171 | 0.17366 | 6220 | 179 | 10.2517 | 10.2513 | 367170 | 607 |  | 35817 | 167 | 1.0068 | 1.0199
BHVO1 |  | 7.23 | 0.055421 | 1971 | 69 | 13.8 | 0.67034 | 23840 | 169 | 49.94 | 5.7477 | 204411 | 454 | 0.273 | 0.038691 | 1376 | 80 | 0.52 | 0.34976 | 12439 | 178 | 11.4 | 10.5621 | 375629 | 607 | 2.71 | 2.5329 | 90081 | 324 | 289 | 0.10168 | 3616 | 186 | 0.168 | 0.17602 | 6260 | 181 | 11.1176 | 11.2172 | 398929 | 632 |  | 35564 | 209 | 1.0139 | 1.0113
BHVO1 |  | 7.23 | 0.059436 | 2125 | 71 | 13.8 | 0.66249 | 23686 | 168 | 49.94 | 5.7257 | 204712 | 455 | 0.273 | 0.039689 | 1419 | 81 | 0.52 | 0.34943 | 12493 | 176 | 11.4 | 10.4802 | 374698 | 606 | 2.71 | 2.4851 | 88848 | 323 | 289 | 0.11339 | 4054 | 187 | 0.168 | 0.17218 | 6156 | 181 | 11.1176 | 11.1499 | 398641 | 633 |  | 35753 | 209 | 1.0086 | 1.0282
BHVO1 |  | 7.23 | 0.058373 | 2089 | 71 | 13.8 | 0.66711 | 23874 | 169 | 49.94 | 5.7341 | 205205 | 455 | 0.273 | 0.043871 | 1570 | 81 | 0.52 | 0.35639 | 12754 | 177 | 11.4 | 10.5021 | 375839 | 606 | 2.71 | 2.5011 | 89506 | 323 | 289 | 0.098052 | 3509 | 187 | 0.168 | 0.1735 | 6209 | 181 | 11.1176 | 11.1387 | 398621 | 633 |  | 35787 | 210 | 1.0076 | 1.0351
BHVO1 |  | 7.23 | 0.060676 | 2177 | 70 | 13.8 | 0.6719 | 24107 | 169 | 49.94 | 5.7174 | 205133 | 455 | 0.273 | 0.041027 | 1472 | 81 | 0.52 | 0.35405 | 12703 | 177 | 11.4 | 10.4255 | 374058 | 606 | 2.71 | 2.4939 | 89479 | 323 | 289 | 0.10566 | 3791 | 187 | 0.168 | 0.17096 | 6134 | 182 | 11.1176 | 11.1076 | 398530 | 633 |  | 35879 | 210 | 1.005 | 1.047
BHVO1 |  | 7.23 | 0.05833 | 2070 | 70 | 13.8 | 0.66825 | 23715 | 168 | 49.94 | 5.7599 | 204409 | 454 | 0.273 | 0.043367 | 1539 | 82 | 0.52 | 0.35756 | 12689 | 176 | 11.4 | 10.5373 | 373947 | 605 | 2.71 | 2.5082 | 89010 | 323 | 289 | 0.10451 | 3709 | 186 | 0.168 | 0.174 | 6175 | 181 | 11.1176 | 11.1364 | 395208 | 630 |  | 35488 | 209 | 1.0161 | 1.0192
BCR1 |  | 3.48 | 0.018658 | 688 | 56 | 13.65 | 0.67721 | 24972 | 172 | 54.11 | 6.427 | 236996 | 487 | 0.36 | 0.060529 | 2232 | 83 | 1.69 | 1.3123 | 48391 | 266 | 6.95 | 6.3995 | 235983 | 494 | 2.24 | 2.1794 | 80366 | 312 | 16 | 0.018875 | 696 | 167 | 0.18 | 0.19761 | 7287 | 189 | 12.1105 | 12.6765 | 467447 | 681 |  | 36875 | 172 | 0.97789 | 0.98698
BCR1 |  | 3.48 | 0.02108 | 773 | 56 | 13.65 | 0.69258 | 25397 | 172 | 54.11 | 6.493 | 238099 | 488 | 0.36 | 0.063076 | 2313 | 83 | 1.69 | 1.3232 | 48522 | 266 | 6.95 | 6.3631 | 233335 | 492 | 2.24 | 2.1923 | 80390 | 311 | 16 | 0.033515 | 1229 | 167 | 0.18 | 0.2213 | 8115 | 189 | 12.1105 | 12.7709 | 468310 | 681 |  | 36670 | 171 | 0.98335 | 0.97397
BCR1 |  | 3.48 | 0.018242 | 672 | 55 | 13.65 | 0.68038 | 25064 | 172 | 54.11 | 6.4621 | 238052 | 489 | 0.36 | 0.062381 | 2298 | 83 | 1.69 | 1.3075 | 48166 | 266 | 6.95 | 6.3755 | 234862 | 493 | 2.24 | 2.1928 | 80778 | 311 | 16 | 0.03814 | 1405 | 167 | 0.18 | 0.2021 | 7445 | 189 | 12.1105 | 12.725 | 468763 | 682 |  | 36838 | 172 | 0.97887 | 0.94789
BCR1 |  | 3.48 | 0.020069 | 749 | 57 | 13.65 | 0.66818 | 24938 | 172 | 54.11 | 6.3527 | 237094 | 488 | 0.36 | 0.060393 | 2254 | 83 | 1.69 | 1.2973 | 48417 | 267 | 6.95 | 6.2943 | 234917 | 493 | 2.24 | 2.1647 | 80791 | 312 | 16 | 0.033707 | 1258 | 167 | 0.18 | 0.20042 | 7480 | 189 | 12.1105 | 12.586 | 469736 | 682 |  | 37322 | 172 | 0.96618 | 0.96522
BCR1 |  | 3.48 | 0.019724 | 701 | 56 | 13.65 | 0.68663 | 24403 | 170 | 54.11 | 6.523 | 231829 | 483 | 0.36 | 0.062183 | 2210 | 82 | 1.69 | 1.326 | 47125 | 262 | 6.95 | 6.3815 | 226800 | 485 | 2.24 | 2.2138 | 78680 | 308 | 16 | 0.034918 | 1241 | 172 | 0.18 | 0.20473 | 7276 | 188 | 12.1105 | 12.7724 | 453932 | 673 |  | 35540 | 209 | 1.0146 | 1.022
Method #1: Sensitivity Counts/Unit Concentation
Element |  | MgO |  |  |  | Al2O3 |  |  |  | SiO2 |  |  |  | P2O5 |  |  |  | K2O |  |  |  | CaO |  |  |  | TiO2 |  |  |  | Cr ppm |  |  |  | MnO |  |  |  | FeO |  |  |  |  | Rh-La Coh Area
|  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD
BIR1 |  | 323.657 | 3133 | 2.527% |  | 1804.2606 | 27695.4 | 0.6534% |  | 4011.9112 | 191649 | 0.2315% |  | 5704.3478 | 262.4 | 25.09% |  |  | -232 | -29.59% |  | 33583.3535 | 444643.6 | 0.1474% |  | 31357.7083 | 30103.4 | 0.7267% |  | 13.5105 | 5161 | 3.548% |  | 38394.152 | 6565.4 | 2.878% |  | 36602.9113 | 375241.4 | 0.1653% |  |  | 35835.6 | 0.4663%
BHVO1 |  | 288.5754 | 2086.4 | 3.365% |  | 1727.8551 | 23844.4 | 0.7071% |  | 4100.4005 | 204774 | 0.222% |  | 5403.663 | 1475.2 | 5.491% |  | 24260.7692 | 12615.6 | 1.401% |  | 32880.193 | 374834.2 | 0.1617% |  | 32983.321 | 89384.8 | 0.3616% |  | 12.9266 | 3735.8 | 4.995% |  | 36826.1905 | 6186.8 | 2.929% |  | 35797.9171 | 397985.8 | 0.1588% |  |  | 35694.2 | 0.5866%
BCR1 |  | 205.9195 | 716.6 | 7.815% |  | 1828.1905 | 24954.8 | 0.6876% |  | 4369.1369 | 236414 | 0.206% |  | 6281.6667 | 2261.4 | 3.661% |  | 28475.858 | 48124.2 | 0.5515% |  | 33550.9928 | 233179.4 | 0.2107% |  | 35804.0179 | 80201 | 0.3875% |  | 72.8625 | 1165.8 | 14.41% |  | 41781.1111 | 7520.6 | 2.51% |  | 38449.2375 | 465637.6 | 0.146% |  |  | 36649 | 0.489%
Average Total Counts | Avg | 272.7173 |  |  |  | 1786.7687 |  |  |  | 4160.4829 |  |  |  | 5796.5592 |  |  |  | 26368.3136 |  |  |  | 33338.1798 |  |  |  | 33381.6824 |  |  |  | 33.0999 |  |  |  | 39000.4845 |  |  |  | 36950.022 |  |  |  |  | 36059.6
/wt% or ppm | stdev | 60.4495 |  |  |  | 52.4049 |  |  |  | 186.0376 |  |  |  | 446.206 |  |  |  | 2980.5178 |  |  |  | 396.9581 |  |  |  | 2249.7635 |  |  |  | 34.4367 |  |  |  | 2532.4966 |  |  |  | 1359.3157 |  |  |  |  | 515.3084
| %std | 22.17% |  |  |  | 2.933% |  |  |  | 4.472% |  |  |  | 7.698% |  |  |  | 11.3% |  |  |  | 1.191% |  |  |  | 6.74% |  |  |  | 104% |  |  |  | 6.494% |  |  |  | 3.679% |  |  |  |  | 1.429%
Method #2: |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept
Slope Correction |  | 0.69048 | 1.7623 |  |  | 0.77559 | 3.1957 |  |  | 0.6559 | 17.3758 |  |  | 0.93951 | 0.01132 |  |  | 0.91511 | 0.053109 |  |  | 0.9661 | 0.33377 |  |  | 0.94492 | 0.087482 |  |  | 3.0047 | -77.0805 |  |  | 0.34476 | 0.11331 |  |  | 0.81156 | 2.0974
|  | 0.013477 | 0.11009 |  |  | 0.057786 | 0.82669 |  |  | 0.020569 | 1.0453 |  |  | 0.032397 | 0.0086885 |  |  | 0.0072025 | 0.0077342 |  |  | 0.011034 | 0.12029 |  |  | 0.022455 | 0.04804 |  |  | 0.13616 | 15.5129 |  |  | 0.043 | 0.0074676 |  |  | 0.046384 | 0.5197
|  | 0.99507 | 0.19225 |  |  | 0.93269 | 0.21416 |  |  | 0.98738 | 0.31751 |  |  | 0.98478 | 0.017543 |  |  | 0.9992 | 0.021251 |  |  | 0.99831 | 0.1167 |  |  | 0.99271 | 0.067813 |  |  | 0.974 | 26.9045 |  |  | 0.83179 | 0.0022464 |  |  | 0.95927 | 0.16464
|  | 2625.0846 | 13 |  |  | 180.1477 | 13 |  |  | 1016.8634 | 13 |  |  | 840.9768 | 13 |  |  | 16143.0758 | 13 |  |  | 7666.5191 | 13 |  |  | 1770.8169 | 13 |  |  | 486.9506 | 13 |  |  | 64.2821 | 13 |  |  | 306.1373 | 13
|  | 97.0278 | 0.4805 |  |  | 8.2621 | 0.59622 |  |  | 102.5118 | 1.3106 |  |  | 0.25882 | 0.0040009 |  |  | 7.29 | 0.0058706 |  |  | 104.41 | 0.17705 |  |  | 8.1432 | 0.059781 |  |  | 352479.9301 | 9410.0699 |  |  | 0.0003244 | 6.5604e-05 |  |  | 8.2986 | 0.3524
Method #3: Traditional |  | MgO |  |  |  | Al2O3 |  |  |  | SiO2 |  |  |  | P2O5 |  |  |  | K2O |  |  |  | CaO |  |  |  | TiO2 |  |  |  | Cr ppm |  |  |  | MnO |  |  |  | FeO |  |  |  |  | Rh-La Coh Area
Rh-tube line Corrected Counts |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept
|  | 91.2974 | 1.7623 |  |  | 15.6526 | 3.1957 |  |  | 5.6848 | 17.3758 |  |  | 5.8446 | 0.01132 |  |  | 1.2515 | 0.053109 |  |  | 1.045 | 0.33377 |  |  | 1.0207 | 0.087482 |  |  | 3273.3383 | -77.0805 |  |  | 0.31876 | 0.11331 |  |  | 0.79201 | 2.0974
|  | 1.7819 | 0.11009 |  |  | 1.1662 | 0.82669 |  |  | 0.17827 | 1.0453 |  |  | 0.20154 | 0.0086885 |  |  | 0.0098497 | 0.0077342 |  |  | 0.011934 | 0.12029 |  |  | 0.024256 | 0.04804 |  |  | 148.3366 | 15.5129 |  |  | 0.039758 | 0.0074676 |  |  | 0.045266 | 0.5197
|  | 0.99507 | 0.19225 |  |  | 0.93269 | 0.21416 |  |  | 0.98738 | 0.31751 |  |  | 0.98478 | 0.017543 |  |  | 0.9992 | 0.021251 |  |  | 0.99831 | 0.1167 |  |  | 0.99271 | 0.067813 |  |  | 0.974 | 26.9045 |  |  | 0.83179 | 0.0022464 |  |  | 0.95927 | 0.16464
|  | 2625.0846 | 13 |  |  | 180.1477 | 13 |  |  | 1016.8634 | 13 |  |  | 840.9768 | 13 |  |  | 16143.0758 | 13 |  |  | 7666.5191 | 13 |  |  | 1770.8169 | 13 |  |  | 486.9506 | 13 |  |  | 64.2821 | 13 |  |  | 306.1373 | 13
|  | 97.0278 | 0.4805 |  |  | 8.2621 | 0.59622 |  |  | 102.5118 | 1.3106 |  |  | 0.25882 | 0.0040009 |  |  | 7.29 | 0.0058706 |  |  | 104.41 | 0.17705 |  |  | 8.1432 | 0.059781 |  |  | 352479.9301 | 9410.0699 |  |  | 0.0003244 | 6.5604e-05 |  |  | 8.2986 | 0.3524
Standards as Unknowns |  | MgO | MgO calc | MgO slope | Trad | Al2O3 | Al2O3 calc | Al slope | Trad | SiO2 | SiO2 calc | Si slope | Trad | P2O5 | P2O5 calc | P slope | Trad | K2O | K2O calc | K slope | Trad | CaO | CaO calc | Ca slope | Trad | TiO2 | TiO2 calc | Ti slope | Trad | Cr ppm | Cr calc | Cr slope | Trad | MnO | MnO calc | MnO slope | Trad | FeO | FeO calc | FeO slope | Trad | Rh-La Coh Area | Rh-La-Coh AreaStd
BIR1 |  | 9.68 | 11.4639 | 9.6778 | 9.6778 | 15.35 | 15.5819 | 15.2809 | 15.2809 | 47.77 | 46.7065 | 48.0107 | 48.0107 | 0.046 | 0.045758 | 0.054311 | 0.054311 | 0.027 | -0.0051251 | 0.048419 | 0.048419 | 13.24 | 13.4922 | 13.3686 | 13.3686 | 0.96 | 0.91962 | 0.95644 | 0.95644 | 382 | 156.4582 | 393.0241 | 393.0241 | 0.171 | 0.17701 | 0.17433 | 0.17433 | 10.2517 | 10.2816 | 10.4415 | 10.4415 | 35755 | 208
BIR1 |  | 9.68 | 11.6664 | 9.8177 | 9.8177 | 15.35 | 15.6005 | 15.2954 | 15.2954 | 47.77 | 46.0729 | 47.5951 | 47.5951 | 0.046 | 0.040964 | 0.049806 | 0.049806 | 0.027 | -0.0042377 | 0.049231 | 0.049231 | 13.24 | 13.3558 | 13.2369 | 13.2369 | 0.96 | 0.89746 | 0.93551 | 0.93551 | 382 | 152.3069 | 380.5511 | 380.5511 | 0.171 | 0.17027 | 0.17201 | 0.17201 | 10.2517 | 10.1904 | 10.3675 | 10.3675 | 36143 | 168
BIR1 |  | 9.68 | 11.9462 | 10.0109 | 10.0109 | 15.35 | 15.7818 | 15.436 | 15.436 | 47.77 | 46.8024 | 48.0736 | 48.0736 | 0.046 | 0.04948 | 0.057807 | 0.057807 | 0.027 | -0.0090771 | 0.044803 | 0.044803 | 13.24 | 13.5554 | 13.4297 | 13.4297 | 0.96 | 0.91713 | 0.95409 | 0.95409 | 382 | 163.7209 | 414.8461 | 414.8461 | 0.171 | 0.16865 | 0.17145 | 0.17145 | 10.2517 | 10.2994 | 10.456 | 10.456 | 35706 | 208
BIR1 |  | 9.68 | 11.6186 | 9.7846 | 9.7846 | 15.35 | 15.844 | 15.4842 | 15.4842 | 47.77 | 46.8156 | 48.0822 | 48.0822 | 0.046 | 0.044538 | 0.053164 | 0.053164 | 0.027 | -0.0109 | 0.043135 | 0.043135 | 13.24 | 13.5252 | 13.4006 | 13.4006 | 0.96 | 0.915 | 0.95209 | 0.95209 | 382 | 162.9085 | 412.4052 | 412.4052 | 0.171 | 0.17048 | 0.17208 | 0.17208 | 10.2517 | 10.3194 | 10.4722 | 10.4722 | 35757 | 168
BIR1 |  | 9.68 | 11.1044 | 9.4296 | 9.4296 | 15.35 | 15.1785 | 14.9681 | 14.9681 | 47.77 | 45.3683 | 47.1329 | 47.1329 | 0.046 | 0.047069 | 0.055542 | 0.055542 | 0.027 | -0.014967 | 0.039413 | 0.039413 | 13.24 | 13.1763 | 13.0635 | 13.0635 | 0.96 | 0.88811 | 0.92667 | 0.92667 | 382 | 149.1612 | 371.0991 | 371.0991 | 0.171 | 0.16057 | 0.16867 | 0.16867 | 10.2517 | 10.0042 | 10.2164 | 10.2164 | 35817 | 167
BHVO1 |  | 7.23 | 7.328 | 6.8221 | 6.8221 | 13.8 | 13.5285 | 13.6883 | 13.6883 | 49.94 | 49.8162 | 50.0504 | 50.0504 | 0.273 | 0.24069 | 0.23745 | 0.23745 | 0.52 | 0.47831 | 0.49082 | 0.49082 | 11.4 | 11.4242 | 11.3708 | 11.3708 | 2.71 | 2.7361 | 2.6729 | 2.6729 | 289 | 110.7675 | 255.7389 | 255.7389 | 0.168 | 0.16275 | 0.16942 | 0.16942 | 11.1176 | 10.9469 | 10.9815 | 10.9815 | 35564 | 209
BHVO1 |  | 7.23 | 7.8588 | 7.1886 | 7.1886 | 13.8 | 13.37 | 13.5654 | 13.5654 | 49.94 | 49.6259 | 49.9255 | 49.9255 | 0.273 | 0.2469 | 0.24329 | 0.24329 | 0.52 | 0.47785 | 0.4904 | 0.4904 | 11.4 | 11.3357 | 11.2852 | 11.2852 | 2.71 | 2.6844 | 2.624 | 2.624 | 289 | 123.5281 | 294.0803 | 294.0803 | 0.168 | 0.1592 | 0.16819 | 0.16819 | 11.1176 | 10.8812 | 10.9281 | 10.9281 | 35753 | 209
BHVO1 |  | 7.23 | 7.7183 | 7.0916 | 7.0916 | 13.8 | 13.4633 | 13.6378 | 13.6378 | 49.94 | 49.6981 | 49.9729 | 49.9729 | 0.273 | 0.27291 | 0.26773 | 0.26773 | 0.52 | 0.48737 | 0.49911 | 0.49911 | 11.4 | 11.3594 | 11.3081 | 11.3081 | 2.71 | 2.7017 | 2.6404 | 2.6404 | 289 | 106.82 | 243.878 | 243.878 | 0.168 | 0.16042 | 0.16861 | 0.16861 | 11.1176 | 10.8703 | 10.9193 | 10.9193 | 35787 | 210
BHVO1 |  | 7.23 | 8.0228 | 7.3018 | 7.3018 | 13.8 | 13.5599 | 13.7127 | 13.7127 | 49.94 | 49.5533 | 49.8779 | 49.8779 | 0.273 | 0.25522 | 0.2511 | 0.2511 | 0.52 | 0.48418 | 0.49619 | 0.49619 | 11.4 | 11.2766 | 11.2281 | 11.2281 | 2.71 | 2.694 | 2.6331 | 2.6331 | 289 | 115.1087 | 268.7827 | 268.7827 | 0.168 | 0.15807 | 0.16781 | 0.16781 | 11.1176 | 10.8399 | 10.8947 | 10.8947 | 35879 | 210
BHVO1 |  | 7.23 | 7.7125 | 7.0876 | 7.0876 | 13.8 | 13.4863 | 13.6556 | 13.6556 | 49.94 | 49.9224 | 50.12 | 50.12 | 0.273 | 0.26978 | 0.26478 | 0.26478 | 0.52 | 0.48897 | 0.50058 | 0.50058 | 11.4 | 11.3974 | 11.3449 | 11.3449 | 2.71 | 2.7094 | 2.6476 | 2.6476 | 289 | 113.8597 | 265.0298 | 265.0298 | 0.168 | 0.16088 | 0.16877 | 0.16877 | 11.1176 | 10.868 | 10.9175 | 10.9175 | 35488 | 209
BCR1 |  | 3.48 | 2.467 | 3.4657 | 3.4657 | 13.65 | 13.667 | 13.7958 | 13.7958 | 54.11 | 55.704 | 53.9121 | 53.9121 | 0.36 | 0.37654 | 0.36509 | 0.36509 | 1.69 | 1.7946 | 1.6954 | 1.6954 | 6.95 | 6.9219 | 7.0211 | 7.0211 | 2.24 | 2.3543 | 2.3121 | 2.3121 | 16 | 20.5623 | -15.2977 | -15.2977 | 0.18 | 0.18271 | 0.1763 | 0.1763 | 12.1105 | 12.371 | 12.1373 | 12.1373 | 36875 | 172
BCR1 |  | 3.48 | 2.7873 | 3.6868 | 3.6868 | 13.65 | 13.9773 | 14.0364 | 14.0364 | 54.11 | 56.2761 | 54.2874 | 54.2874 | 0.36 | 0.39239 | 0.37997 | 0.37997 | 1.69 | 1.8095 | 1.709 | 1.709 | 6.95 | 6.8825 | 6.983 | 6.983 | 2.24 | 2.3681 | 2.3252 | 2.3252 | 16 | 36.512 | 32.6258 | 32.6258 | 0.18 | 0.20461 | 0.18385 | 0.18385 | 12.1105 | 12.4632 | 12.212 | 12.212 | 36670 | 171
BCR1 |  | 3.48 | 2.412 | 3.4277 | 3.4277 | 13.65 | 13.7311 | 13.8455 | 13.8455 | 54.11 | 56.0084 | 54.1118 | 54.1118 | 0.36 | 0.38807 | 0.37591 | 0.37591 | 1.69 | 1.7881 | 1.6894 | 1.6894 | 6.95 | 6.896 | 6.996 | 6.996 | 2.24 | 2.3687 | 2.3257 | 2.3257 | 16 | 41.5504 | 47.7645 | 47.7645 | 0.18 | 0.18686 | 0.17773 | 0.17773 | 12.1105 | 12.4183 | 12.1756 | 12.1756 | 36838 | 172
BCR1 |  | 3.48 | 2.6535 | 3.5945 | 3.5945 | 13.65 | 13.4849 | 13.6546 | 13.6546 | 54.11 | 55.0596 | 53.4895 | 53.4895 | 0.36 | 0.3757 | 0.36429 | 0.36429 | 1.69 | 1.7741 | 1.6766 | 1.6766 | 6.95 | 6.8081 | 6.9111 | 6.9111 | 2.24 | 2.3384 | 2.297 | 2.297 | 16 | 36.7206 | 33.2528 | 33.2528 | 0.18 | 0.18531 | 0.17719 | 0.17719 | 12.1105 | 12.2827 | 12.0656 | 12.0656 | 37322 | 172
BCR1 |  | 3.48 | 2.608 | 3.563 | 3.563 | 13.65 | 13.8573 | 13.9433 | 13.9433 | 54.11 | 56.5363 | 54.4581 | 54.4581 | 0.36 | 0.38683 | 0.37476 | 0.37476 | 1.69 | 1.8133 | 1.7125 | 1.7125 | 6.95 | 6.9025 | 7.0023 | 7.0023 | 2.24 | 2.3914 | 2.3472 | 2.3472 | 16 | 38.0407 | 37.2192 | 37.2192 | 0.18 | 0.18929 | 0.17857 | 0.17857 | 12.1105 | 12.4646 | 12.2132 | 12.2132 | 35540 | 209
|  | MgO | Mg calc | Mg Slope | MgO trad | Al2O3 | Al calc | Al Slope | Al2O3 Trad | SiO2 | SiO2 calc | SiO2 slope | SiO2 trad | P2O5 | P2O5 calc | P2O5 slope | P2O5 trad | K2O | K2O calc | K2O slope | K2O trad | CaO | CaO calc | CaO slope | CaO trad | TiO2 | TiO2 calc | TiO2 slope | TiO2 trad | Cr ppm | Cr calc | Cr slope | Cr trad | MnO | MnO calc | MnO slope | MnO trad | FeO | FeO calc | FeO* slope | FeO* trad
BIR-1 |  | 9.68 | 11.5599 | 9.7441 | 9.7441 | 15.35 | 15.5973 | 15.2929 | 15.2929 | 47.77 | 46.3531 | 47.7789 | 47.7789 | 0.046 | 0.045562 | 0.054126 | 0.054126 | 0.027 | -0.0088614 | 0.045 | 0.045 | 13.24 | 13.421 | 13.2999 | 13.2999 | 0.96 | 0.90746 | 0.94496 | 0.94496 | 382 | 156.9111 | 394.3851 | 394.3851 | 0.171 | 0.16939 | 0.17171 | 0.17171 | 10.2517 | 10.219 | 10.3907 | 10.3907
BHVO-1 |  | 7.23 | 7.7281 | 7.0983 | 7.0983 | 13.8 | 13.4816 | 13.652 | 13.652 | 49.94 | 49.7232 | 49.9893 | 49.9893 | 0.273 | 0.2571 | 0.25287 | 0.25287 | 0.52 | 0.48334 | 0.49542 | 0.49542 | 11.4 | 11.3587 | 11.3074 | 11.3074 | 2.71 | 2.7051 | 2.6436 | 2.6436 | 289 | 114.0168 | 265.5019 | 265.5019 | 0.168 | 0.16026 | 0.16856 | 0.16856 | 11.1176 | 10.8813 | 10.9282 | 10.9282
BCR-1 |  | 3.48 | 2.5856 | 3.5475 | 3.5475 | 13.65 | 13.7435 | 13.8551 | 13.8551 | 54.11 | 55.9169 | 54.0518 | 54.0518 | 0.36 | 0.38391 | 0.372 | 0.372 | 1.69 | 1.7959 | 1.6966 | 1.6966 | 6.95 | 6.8822 | 6.9827 | 6.9827 | 2.24 | 2.3642 | 2.3214 | 2.3214 | 16 | 34.6772 | 27.1129 | 27.1129 | 0.18 | 0.18976 | 0.17873 | 0.17873 | 12.1105 | 12.4 | 12.1607 | 12.1607

## STable3B-30 KV XRF2
USGS Standards | Calibration Data | K2O | Counts /Rh | K -Ka Area | K -Ka AreaStd | CaO | Counts /Rh | Ca-Ka Area | Ca-Ka AreaStd | TiO2 | Counts /Rh | Ti-Ka Area | Ti-Ka AreaStd | MnO | Counts /Rh | Mn-Ka Area | Mn-Ka AreaStd | FeO | Counts /Rh | Fe-Ka Area | Fe-Ka AreaStd | Co ppm | Counts /Rh | Co-Ka Area | Co-Ka AreaStd | Ni ppm | Counts /Rh | Ni-Ka Area | Ni-Ka AreaStd | Cu ppm | Counts /Rh | Cu-Ka Area | Cu-Ka AreaStd | Zn ppm | Counts /Rh | Zn-Ka Area | Zn-Ka AreaStd | Rb ppm | Counts /Rh | Rb-Ka Area | Rb-Ka AreaStd | Sr ppm | Counts /Rh | Sr-Ka Area | Sr-Ka AreaStd | Y ppm | Counts /Rh | Y -Ka Area | Y -Ka AreaStd | Zr ppm | Counts /Rh | Zr-Ka Area | Zr-Ka AreaStd | Nb ppm | Counts /Rh | Nb-Ka Area | Nb-Ka AreaStd | Rh-coh/ Rh-Inc | Rh-Ka-Coh Area | Rh-Ka-Coh AreaStd | Rh-Ka-Inc Area | Rh-Ka-Inc AreaStd | RH-Co ratio | Rh inc ratio | Rayleigh/ Compton
BIR1 | BIR1 | 0.027 | 79.4232 | 73 | 30 | 13.24 | 27315.0633 | 25106 | 154 | 0.96 | 3581.6613 | 3292 | 65 | 0.171 | 2964.7713 | 2725 | 66 | 10.2517 | 270081.4144 | 248239 | 477 | 51.4 | 3579.4853 | 3290 | 108 | 166 | 4715.3463 | 4334 | 77 | 126 | 1212.0203 | 1114 | 46 | 71 | 1193.5244 | 1097 | 44 | 0.27 | 484.1553 | 445 | 63 | 108 | 5829.4475 | 5358 | 90 | 16 | 775.7365 | 713 | 73 | 22 | 1387.1866 | 1275 | 101 | 2 | 319.8689 | 294 | 118 | 0.91913 | 31106 | 306 | 33843 | 505 | 1.0808 | 0.9934 | 1.088
BIR1 | BIR1 | 0.027 | 38.6481 | 34 | 28 | 13.24 | 28456.3509 | 25034 | 154 | 0.96 | 3729.5393 | 3281 | 64 | 0.171 | 3120.2638 | 2745 | 66 | 10.2517 | 281562.5997 | 247700 | 477 | 51.4 | 3763.6406 | 3311 | 107 | 166 | 4919.6727 | 4328 | 76 | 126 | 1361.7763 | 1198 | 46 | 71 | 1234.465 | 1086 | 44 | 0.27 | 686.5717 | 604 | 63 | 108 | 6242.801 | 5492 | 90 | 16 | 840.0273 | 739 | 72 | 22 | 1353.8194 | 1191 | 101 | 2 | 230.7517 | 203 | 119 | 0.87973 | 30481 | 306 | 34648 | 506 | 1.103 | 0.97032 | 1.1367
BIR1 | BIR1 | 0.027 | 47.2577 | 43 | 29 | 13.24 | 27308.3348 | 24848 | 153 | 0.96 | 3505.8591 | 3190 | 64 | 0.171 | 2947.5593 | 2682 | 66 | 10.2517 | 271568.905 | 247102 | 477 | 51.4 | 3555.3148 | 3235 | 108 | 166 | 4698.2909 | 4275 | 77 | 126 | 1332.0067 | 1212 | 47 | 71 | 1283.65 | 1168 | 44 | 0.27 | 685.7856 | 624 | 64 | 108 | 6095.1394 | 5546 | 92 | 16 | 797.8852 | 726 | 73 | 22 | 1375.9673 | 1252 | 101 | 2 | 249.4765 | 227 | 117 | 0.90991 | 30571 | 306 | 33598 | 505 | 1.0997 | 1.0006 | 1.099
BIR1 | BIR1 | 0.027 | 58.3206 | 53 | 29 | 13.24 | 27396.3844 | 24897 | 153 | 0.96 | 3619.1794 | 3289 | 64 | 0.171 | 3117.402 | 2833 | 66 | 10.2517 | 271501.1735 | 246732 | 476 | 51.4 | 3701.7085 | 3364 | 109 | 166 | 4705.2633 | 4276 | 76 | 126 | 1294.0574 | 1176 | 46 | 71 | 1294.0574 | 1176 | 45 | 0.27 | 628.3221 | 571 | 64 | 108 | 6222.6997 | 5655 | 92 | 16 | 941.933 | 856 | 73 | 22 | 1538.3438 | 1398 | 101 | 2 | 182.6646 | 166 | 118 | 0.90877 | 30850 | 306 | 33947 | 508 | 1.0898 | 0.99036 | 1.1004
BIR1 | BIR1 | 0.027 | 72.7164 | 68 | 29 | 13.24 | 26160.8056 | 24464 | 152 | 0.96 | 3365.2737 | 3147 | 63 | 0.171 | 2868.0216 | 2682 | 65 | 10.2517 | 259862.8614 | 243008 | 473 | 51.4 | 3316.0831 | 3101 | 107 | 166 | 4380.0956 | 4096 | 75 | 126 | 1249.0116 | 1168 | 46 | 71 | 1210.5147 | 1132 | 43 | 0.27 | 509.015 | 476 | 63 | 108 | 5630.1766 | 5265 | 90 | 16 | 829.8228 | 776 | 71 | 22 | 1297.1328 | 1213 | 100 | 2 | 144.3635 | 135 | 118 | 0.93514 | 30825 | 305 | 32963 | 503 | 1.0907 | 1.0199 | 1.0694
BHVO1 | BHVO1 | 0.52 | 617.6822 | 574 | 40 | 11.4 | 22743.4027 | 21135 | 144 | 2.71 | 10445.7161 | 9707 | 101 | 0.168 | 2789.2548 | 2592 | 66 | 11.1176 | 286490.471 | 266230 | 494 | 45 | 3770.6592 | 3504 | 111 | 121 | 4431.3855 | 4118 | 75 | 136 | 1041.6661 | 968 | 45 | 105 | 1468.8784 | 1365 | 47 | 11 | 1079.3297 | 1003 | 52 | 403 | 20089.7367 | 18669 | 148 | 27.6 | 1290.2456 | 1199 | 78 | 179 | 10645.871 | 9893 | 148 | 19 | 1132.0587 | 1052 | 121 | 0.92928 | 30893 | 303 | 33244 | 497 | 1.0883 | 1.0113 | 1.0761
BHVO1 | BHVO1 | 0.52 | 601.3682 | 569 | 40 | 11.4 | 22266.477 | 21068 | 143 | 2.71 | 10042.532 | 9502 | 101 | 0.168 | 2766.9279 | 2618 | 66 | 11.1176 | 281817.6299 | 266649 | 495 | 45 | 3852.3499 | 3645 | 112 | 121 | 4488.5954 | 4247 | 76 | 136 | 1074.8532 | 1017 | 46 | 105 | 1555.7364 | 1472 | 48 | 11 | 924.7754 | 875 | 50 | 403 | 19592.5551 | 18538 | 148 | 27.6 | 1271.434 | 1203 | 77 | 179 | 10481.1398 | 9917 | 148 | 19 | 1004.0418 | 950 | 122 | 0.94618 | 30939 | 302 | 32699 | 494 | 1.0866 | 1.0282 | 1.0569
BHVO1 | BHVO1 | 0.52 | 586.4948 | 566 | 40 | 11.4 | 22054.692 | 21284 | 144 | 2.71 | 9890.6237 | 9545 | 101 | 0.168 | 2631.9732 | 2540 | 66 | 11.1176 | 276618.3111 | 266952 | 495 | 45 | 3754.1885 | 3623 | 113 | 121 | 4071.2688 | 3929 | 74 | 136 | 978.1822 | 944 | 45 | 105 | 1553.2787 | 1499 | 48 | 11 | 1033.1013 | 997 | 51 | 403 | 19322.2064 | 18647 | 149 | 27.6 | 1220.6553 | 1178 | 79 | 179 | 10590.0654 | 10220 | 148 | 19 | 949.1683 | 916 | 121 | 0.96506 | 31345 | 304 | 32480 | 495 | 1.0726 | 1.0351 | 1.0362
BHVO1 | BHVO1 | 0.52 | 595.2508 | 567 | 39 | 11.4 | 22473.6057 | 21407 | 144 | 2.71 | 10065.723 | 9588 | 101 | 0.168 | 2581.5199 | 2459 | 65 | 11.1176 | 279974.7035 | 266687 | 495 | 45 | 3955.7409 | 3768 | 113 | 121 | 4118.4638 | 3923 | 75 | 136 | 1045.6258 | 996 | 45 | 105 | 1563.1896 | 1489 | 48 | 11 | 1067.6721 | 1017 | 52 | 403 | 19743.0106 | 18806 | 149 | 27.6 | 1311.2315 | 1249 | 78 | 179 | 10271.4887 | 9784 | 147 | 19 | 1235.6441 | 1177 | 122 | 0.95254 | 30587 | 303 | 32111 | 493 | 1.0991 | 1.047 | 1.0498
BHVO1 | BHVO1 | 0.52 | 539.5016 | 508 | 40 | 11.4 | 22522.0675 | 21207 | 143 | 2.71 | 10143.2672 | 9551 | 100 | 0.168 | 2605.113 | 2453 | 65 | 11.1176 | 282439.7044 | 265948 | 494 | 45 | 3787.1313 | 3566 | 112 | 121 | 4316.0127 | 4064 | 75 | 136 | 1047.1429 | 986 | 45 | 105 | 1598.3266 | 1505 | 48 | 11 | 1138.4758 | 1072 | 52 | 403 | 20164.4031 | 18987 | 150 | 27.6 | 1464.5132 | 1379 | 78 | 179 | 10543.6453 | 9928 | 147 | 19 | 847.4848 | 798 | 122 | 0.94161 | 31059 | 303 | 32985 | 498 | 1.0824 | 1.0192 | 1.062
BCR1 | BCR1 | 1.69 | 2182.9106 | 2002 | 56 | 6.95 | 14272.8768 | 13090 | 116 | 2.24 | 9795.8384 | 8984 | 99 | 0.18 | 3368.1372 | 3089 | 72 | 12.1105 | 353313.1247 | 324032 | 545 | 37 | 4281.863 | 3927 | 119 | 13 | 3908.9582 | 3585 | 73 | 19 | 264.9587 | 243 | 34 | 129.5 | 1907.0482 | 1749 | 50 | 47.2 | 3255.8296 | 2986 | 71 | 330 | 18299.5944 | 16783 | 143 | 38 | 1769.6623 | 1623 | 85 | 190 | 12243.7077 | 11229 | 153 | 14 | 696.7432 | 639 | 124 | 0.91712 | 31240 | 310 | 34063 | 510 | 1.0762 | 0.98698 | 1.0904
BCR1 | BCR1 | 1.69 | 2141.9155 | 1983 | 56 | 6.95 | 14069.8898 | 13026 | 116 | 2.24 | 9626.1982 | 8912 | 99 | 0.18 | 3249.0579 | 3008 | 71 | 12.1105 | 347853.3439 | 322045 | 544 | 37 | 4237.385 | 3923 | 120 | 13 | 3654.11 | 3383 | 72 | 19 | 224.6689 | 208 | 33 | 129.5 | 1784.3895 | 1652 | 48 | 47.2 | 3166.9674 | 2932 | 72 | 330 | 17924.9057 | 16595 | 143 | 38 | 1774.6683 | 1643 | 85 | 190 | 12141.8418 | 11241 | 152 | 14 | 935.4003 | 866 | 123 | 0.92581 | 31957 | 310 | 34518 | 507 | 1.052 | 0.97397 | 1.0801
BCR1 | BCR1 | 1.69 | 2170.6174 | 1935 | 56 | 6.95 | 14672.7004 | 13080 | 116 | 2.24 | 10164.3225 | 9061 | 99 | 0.18 | 3449.4307 | 3075 | 72 | 12.1105 | 361358.9953 | 322134 | 544 | 37 | 4601.4845 | 4102 | 121 | 13 | 3816.2482 | 3402 | 72 | 19 | 210.892 | 188 | 34 | 129.5 | 1928.3159 | 1719 | 50 | 47.2 | 3297.9923 | 2940 | 70 | 330 | 18479.9744 | 16474 | 142 | 38 | 1780.2428 | 1587 | 85 | 190 | 13003.5124 | 11592 | 154 | 14 | 766.1662 | 683 | 124 | 0.89145 | 31618 | 309 | 35468 | 505 | 1.0633 | 0.94789 | 1.1218
BCR1 | BCR1 | 1.69 | 2338.7802 | 2099 | 56 | 6.95 | 14600.9413 | 13104 | 116 | 2.24 | 10176.3123 | 9133 | 100 | 0.18 | 3302.5939 | 2964 | 71 | 12.1105 | 359922.5618 | 323022 | 544 | 37 | 4355.5464 | 3909 | 120 | 13 | 3771.687 | 3385 | 72 | 19 | 232.8752 | 209 | 34 | 129.5 | 1807.2899 | 1622 | 48 | 47.2 | 3240.1967 | 2908 | 70 | 330 | 18757.0395 | 16834 | 143 | 38 | 2118.1616 | 1901 | 86 | 190 | 12753.5389 | 11446 | 153 | 14 | 772.1652 | 693 | 123 | 0.89748 | 31260 | 309 | 34831 | 507 | 1.0755 | 0.96522 | 1.1142
BCR1 | BCR1 | 1.69 | 2089.0834 | 1981 | 56 | 6.95 | 13061.7805 | 12386 | 113 | 2.24 | 9204.1999 | 8728 | 97 | 0.18 | 3081.4244 | 2922 | 70 | 12.1105 | 329697.6464 | 312640 | 535 | 37 | 4030.5284 | 3822 | 117 | 13 | 3607.6498 | 3421 | 71 | 19 | 286.8403 | 272 | 34 | 129.5 | 1738.9695 | 1649 | 48 | 47.2 | 2938.0042 | 2786 | 69 | 330 | 16928.8521 | 16053 | 140 | 38 | 1580.7855 | 1499 | 84 | 190 | 11651.8337 | 11049 | 152 | 14 | 583.1717 | 553 | 122 | 0.94826 | 31195 | 306 | 32897 | 499 | 1.0777 | 1.022 | 1.0546
Method #1: Sensitivity Counts/Unit Concentation
Element |  | K2O |  |  |  | CaO |  |  |  | TiO2 |  |  |  | MnO |  |  |  | FeO |  |  |  | Co ppm |  |  |  | Ni ppm |  |  |  | Cu ppm |  |  |  | Zn ppm |  |  |  | Rb ppm |  |  |  | Sr ppm |  |  |  | Y ppm |  |  |  | Zr ppm |  |  |  | Nb ppm |  |  |  |  | Rh-Ka Coh Area |  | Rh-Ka Inc Area
|  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD
BIR1 |  | 2007.4074 | 54.2 | 53.51% |  | 1878.3837 | 24869.8 | 0.616% |  | 3374.7917 | 3239.8 | 1.975% |  | 15984.7953 | 2733.4 | 2.407% |  | 24050.3173 | 246556.2 | 0.1931% |  | 63.428 | 3260.2 | 3.307% |  | 25.6735 | 4261.8 | 1.788% |  | 9.3143 | 1173.6 | 3.937% |  | 15.9408 | 1131.8 | 3.888% |  | 2014.8148 | 544 | 11.65% |  | 50.5852 | 5463.2 | 1.662% |  | 47.625 | 762 | 9.501% |  | 57.5364 | 1265.8 | 7.963% |  | 102.5 | 205 | 57.56% |  |  | 30766.6 | 0.9939% | 33799.8 | 1.495%
BHVO1 |  | 1070.7692 | 556.8 | 7.148% |  | 1861.4211 | 21220.2 | 0.6767% |  | 3534.5387 | 9578.6 | 1.052% |  | 15073.8095 | 2532.4 | 2.59% |  | 23970.457 | 266493.2 | 0.1856% |  | 80.4711 | 3621.2 | 3.098% |  | 33.5223 | 4056.2 | 1.849% |  | 7.2221 | 982.2 | 4.602% |  | 13.9619 | 1466 | 3.261% |  | 90.2545 | 992.8 | 5.177% |  | 46.4749 | 18729.4 | 0.7945% |  | 44.9855 | 1241.6 | 6.282% |  | 55.5777 | 9948.4 | 1.484% |  | 51.5053 | 978.6 | 12.43% |  |  | 30964.6 | 0.9785% | 32703.8 | 1.515%
BCR1 |  | 1183.432 | 2000 | 2.8% |  | 1861.4676 | 12937.2 | 0.892% |  | 4001.6071 | 8963.6 | 1.102% |  | 16731.1111 | 3011.6 | 2.364% |  | 26487.4202 | 320774.6 | 0.1691% |  | 106.3946 | 3936.6 | 3.033% |  | 264.2462 | 3435.2 | 2.096% |  | 11.7895 | 224 | 15.09% |  | 12.9591 | 1678.2 | 2.908% |  | 61.661 | 2910.4 | 2.419% |  | 50.1448 | 16547.8 | 0.8593% |  | 43.4368 | 1650.6 | 5.15% |  | 59.5337 | 11311.4 | 1.351% |  | 49.0571 | 686.8 | 17.94% |  |  | 31454 | 0.9818% | 34355.4 | 1.472%
Average Total Counts /wt% or ppm |  | 1127.1006 |  |  |  | 1867.0908 |  |  |  | 3636.9792 |  |  |  | 15929.9053 |  |  |  | 24836.0648 |  |  |  | 83.4312 |  |  |  | 29.5979 |  |  |  | 9.4419 |  |  |  | 14.2873 |  |  |  | 75.9578 |  |  |  | 49.0683 |  |  |  | 45.3491 |  |  |  | 57.5492 |  |  |  | 50.2812 |  |  |  |  | 31061.7333 |  | 33619.6667
|  | 511.3576 |  |  |  | 9.78 |  |  |  | 325.7222 |  |  |  | 830.0131 |  |  |  | 1430.6731 |  |  |  | 21.6357 |  |  |  | 5.55 |  |  |  | 2.2864 |  |  |  | 1.5173 |  |  |  | 20.2187 |  |  |  | 2.2567 |  |  |  | 2.1176 |  |  |  | 1.978 |  |  |  | 1.7311 |  |  |  |  | 353.8444 |  | 840.4056
|  | 45.37% |  |  |  | 0.5238% |  |  |  | 8.956% |  |  |  | 5.21% |  |  |  | 5.76% |  |  |  | 25.93% |  |  |  | 18.75% |  |  |  | 24.22% |  |  |  | 10.62% |  |  |  | 26.62% |  |  |  | 4.599% |  |  |  | 4.67% |  |  |  | 3.437% |  |  |  | 3.443% |  |  |  |  | 1.139% |  | 2.5%
Method #2: | Element | K2O |  |  |  | CaO |  |  |  | TiO2 |  |  |  | MnO |  |  |  | FeO |  |  |  | Co ppm |  |  |  | Ni ppm |  |  |  | Cu ppm |  |  |  | Zn ppm |  |  |  | Rb ppm |  |  |  | Sr ppm |  |  |  | Y ppm |  |  |  | Zr ppm |  |  |  | Nb ppm |  |  |  |  | Rh-Ka Coh Area |  | Rh-Ka Inc Area
Slope Correction |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept
|  | 0.86611 | 0.020258 |  |  | 0.89287 | 0.34798 |  |  | 0.86437 | 0.11059 |  |  | 0.2616 | 0.12391 |  |  | 0.47622 | 5.3896 |  |  | -1.321 | 106.2625 |  |  | 4.2481 | -508.5887 |  |  | 1.0601 | -2.5661 |  |  | 1.3514 | -43.965 |  |  | 1.3424 | -8.9468 |  |  | 0.9908 | -14.9079 |  |  | 0.94901 | -0.37739 |  |  | 0.91327 | 2.0248 |  |  | 0.96734 | -1.2221
|  | 0.018439 | 0.021393 |  |  | 0.033593 | 0.39554 |  |  | 0.056869 | 0.13117 |  |  | 0.039852 | 0.00751 |  |  | 0.059209 | 0.72259 |  |  | 0.1855 | 8.709 |  |  | 0.50661 | 72.9168 |  |  | 0.098255 | 10.0609 |  |  | 0.088455 | 9.6649 |  |  | 0.03863 | 0.99998 |  |  | 0.039926 | 12.9064 |  |  | 0.078042 | 2.3759 |  |  | 0.02964 | 4.8435 |  |  | 0.088741 | 1.3357
|  | 0.99414 | 0.057336 |  |  | 0.98193 | 0.38128 |  |  | 0.94673 | 0.18334 |  |  | 0.76822 | 0.0026369 |  |  | 0.83267 | 0.33369 |  |  | 0.79596 | 2.8583 |  |  | 0.84397 | 27.2418 |  |  | 0.89954 | 18.0289 |  |  | 0.94724 | 5.9184 |  |  | 0.98935 | 2.2258 |  |  | 0.97933 | 19.3755 |  |  | 0.91919 | 2.7439 |  |  | 0.98649 | 9.5799 |  |  | 0.90138 | 2.4064
|  | 2206.3614 | 13 |  |  | 706.4333 | 13 |  |  | 231.0257 | 13 |  |  | 43.0883 | 13 |  |  | 64.6904 | 13 |  |  | 50.7119 | 13 |  |  | 70.3157 | 13 |  |  | 116.4096 | 13 |  |  | 233.4016 | 13 |  |  | 1207.5221 | 13 |  |  | 615.8172 | 13 |  |  | 147.8711 | 13 |  |  | 949.3865 | 13 |  |  | 118.8243 | 13
|  | 7.2531 | 0.042736 |  |  | 102.6971 | 1.8899 |  |  | 7.766 | 0.437 |  |  | 0.00029961 | 9.0393e-05 |  |  | 7.2034 | 1.4476 |  |  | 414.3219 | 106.2114 |  |  | 52182.4813 | 9647.5187 |  |  | 37837.809 | 4225.5244 |  |  | 8175.4758 | 455.3575 |  |  | 5982.2589 | 64.4041 |  |  | 231183.0223 | 4880.311 |  |  | 1113.3229 | 97.8771 |  |  | 87130.2541 | 1193.0792 |  |  | 688.0564 | 75.277
Method #3: Traditional |  | K2O |  |  |  | CaO |  |  |  | TiO2 |  |  |  | MnO |  |  |  | FeO |  |  |  | Co ppm |  |  |  | Ni ppm |  |  |  | Cu ppm |  |  |  | Zn ppm |  |  |  | Rb ppm |  |  |  | Sr ppm |  |  |  | Y ppm |  |  |  | Zr ppm |  |  |  | Nb ppm |  |  |  |  | Rh-Ka Coh Area |  | Rh-Ka Inc Area
Rh-tube line Corrected Counts |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept
|  | 0.00076844 | 0.020258 |  |  | 0.00047822 | 0.34798 |  |  | 0.00023766 | 0.11059 |  |  | 1.6422e-05 | 0.12391 |  |  | 1.9175e-05 | 5.3896 |  |  | -0.015833 | 106.2625 |  |  | 0.14353 | -508.5887 |  |  | 0.11228 | -2.5661 |  |  | 0.094586 | -43.965 |  |  | 0.017673 | -8.9468 |  |  | 0.020192 | -14.9079 |  |  | 0.020927 | -0.37739 |  |  | 0.015869 | 2.0248 |  |  | 0.019239 | -1.2221
|  | 1.636e-05 | 0.021393 |  |  | 1.7992e-05 | 0.39554 |  |  | 1.5636e-05 | 0.13117 |  |  | 2.5017e-06 | 0.00751 |  |  | 2.384e-06 | 0.72259 |  |  | 0.0022234 | 8.709 |  |  | 0.017116 | 72.9168 |  |  | 0.010406 | 10.0609 |  |  | 0.0061912 | 9.6649 |  |  | 0.00050857 | 0.99998 |  |  | 0.00081369 | 12.9064 |  |  | 0.0017209 | 2.3759 |  |  | 0.00051504 | 4.8435 |  |  | 0.0017649 | 1.3357
|  | 0.99414 | 0.057336 |  |  | 0.98193 | 0.38128 |  |  | 0.94673 | 0.18334 |  |  | 0.76822 | 0.0026369 |  |  | 0.83267 | 0.33369 |  |  | 0.79596 | 2.8583 |  |  | 0.84397 | 27.2418 |  |  | 0.89954 | 18.0289 |  |  | 0.94724 | 5.9184 |  |  | 0.98935 | 2.2258 |  |  | 0.97933 | 19.3755 |  |  | 0.91919 | 2.7439 |  |  | 0.98649 | 9.5799 |  |  | 0.90138 | 2.4064
|  | 2206.3614 | 13 |  |  | 706.4333 | 13 |  |  | 231.0257 | 13 |  |  | 43.0883 | 13 |  |  | 64.6904 | 13 |  |  | 50.7119 | 13 |  |  | 70.3157 | 13 |  |  | 116.4096 | 13 |  |  | 233.4016 | 13 |  |  | 1207.5221 | 13 |  |  | 615.8172 | 13 |  |  | 147.8711 | 13 |  |  | 949.3865 | 13 |  |  | 118.8243 | 13
|  | 7.2531 | 0.042736 |  |  | 102.6971 | 1.8899 |  |  | 7.766 | 0.437 |  |  | 0.00029961 | 9.0393e-05 |  |  | 7.2034 | 1.4476 |  |  | 414.3219 | 106.2114 |  |  | 52182.4813 | 9647.5187 |  |  | 37837.809 | 4225.5244 |  |  | 8175.4758 | 455.3575 |  |  | 5982.2589 | 64.4041 |  |  | 231183.0223 | 4880.311 |  |  | 1113.3229 | 97.8771 |  |  | 87130.2541 | 1193.0792 |  |  | 688.0564 | 75.277
Standards as Unknowns | ID | K2O | K2O calc | K2O slope | Trad | CaO | CaO calc | CaO slope | Trad | TiO2 | TiO2 calc | TiO2 slope | Trad | MnO | MnO calc | MnO slope | Trad | FeO | FeO calc | FeO slope | Trad | Co ppm | Co calc | Co slope | Trad | Ni ppm | Ni calc | Ni slope | Trad | Cu ppm | Cu calc | Cu slope | Trad | Zn ppm | Zn calc | Zn slope | Trad | Rb ppm | Rb calc | Rb slope | Trad | Sr ppm | Sr calc | Sr slope | Trad | Y ppm | Y calc | Y slope | Trad | Zr ppm | Zr calc | Zr slope | Trad | Nb ppm | Nb calc | Nb slope | Trad | Rh-Ka-Coh Area | Rh-Ka-Coh AreaStd | Rh-Ka-Inc Area | Rh-Ka-Inc AreaStd
BIR1 |  | 0.027 | 0.070467 | 0.08129 | 0.08129 | 13.24 | 14.6297 | 13.4105 | 13.4105 | 0.96 | 0.98479 | 0.96181 | 0.96181 | 0.171 | 0.18611 | 0.17259 | 0.17259 | 10.2517 | 10.8746 | 10.5682 | 10.5682 | 51.4 | 42.9034 | 49.587 | 49.587 | 166 | 159.3135 | 168.1964 | 168.1964 | 126 | 128.3656 | 133.5153 | 133.5153 | 71 | 83.5376 | 68.9253 | 68.9253 | 0.27 | 6.374 | -0.39043 | -0.39043 | 108 | 118.8027 | 102.8019 | 102.8019 | 16 | 17.1059 | 15.8563 | 15.8563 | 22 | 24.1043 | 24.0387 | 24.0387 | 2 | 6.3616 | 4.9317 | 4.9317 | 31106 | 306 | 33843 | 505
BIR1 |  | 0.027 | 0.03429 | 0.049956 | 0.049956 | 13.24 | 15.241 | 13.9563 | 13.9563 | 0.96 | 1.0254 | 0.99696 | 0.99696 | 0.171 | 0.19587 | 0.17515 | 0.17515 | 10.2517 | 11.3368 | 10.7884 | 10.7884 | 51.4 | 45.1107 | 46.6712 | 46.6712 | 166 | 166.2169 | 197.523 | 197.523 | 126 | 144.2263 | 150.3294 | 150.3294 | 71 | 86.4031 | 72.7977 | 72.7977 | 0.27 | 9.0389 | 3.1868 | 3.1868 | 108 | 127.2267 | 111.1484 | 111.1484 | 16 | 18.5236 | 17.2017 | 17.2017 | 22 | 23.5245 | 23.5091 | 23.5091 | 2 | 4.5892 | 3.2173 | 3.2173 | 30481 | 306 | 34648 | 506
BIR1 |  | 0.027 | 0.041929 | 0.056572 | 0.056572 | 13.24 | 14.6261 | 13.4073 | 13.4073 | 0.96 | 0.96395 | 0.9438 | 0.9438 | 0.171 | 0.18503 | 0.17231 | 0.17231 | 10.2517 | 10.9345 | 10.5968 | 10.5968 | 51.4 | 42.6137 | 49.9697 | 49.9697 | 166 | 158.7373 | 165.7485 | 165.7485 | 126 | 141.0734 | 146.987 | 146.987 | 71 | 89.8457 | 77.4499 | 77.4499 | 0.27 | 9.0285 | 3.1729 | 3.1729 | 108 | 124.2174 | 108.1668 | 108.1668 | 16 | 17.5943 | 16.3198 | 16.3198 | 22 | 23.9094 | 23.8606 | 23.8606 | 2 | 4.9616 | 3.5775 | 3.5775 | 30571 | 306 | 33598 | 505
BIR1 |  | 0.027 | 0.051744 | 0.065074 | 0.065074 | 13.24 | 14.6733 | 13.4494 | 13.4494 | 0.96 | 0.99511 | 0.97073 | 0.97073 | 0.171 | 0.19569 | 0.1751 | 0.1751 | 10.2517 | 10.9317 | 10.5955 | 10.5955 | 51.4 | 44.3684 | 47.6518 | 47.6518 | 166 | 158.9729 | 166.7492 | 166.7492 | 126 | 137.0542 | 142.7261 | 142.7261 | 71 | 90.5741 | 78.4343 | 78.4343 | 0.27 | 8.272 | 2.1574 | 2.1574 | 108 | 126.817 | 110.7425 | 110.7425 | 16 | 20.7707 | 19.3343 | 19.3343 | 22 | 26.7309 | 26.4374 | 26.4374 | 2 | 3.6329 | 2.2921 | 2.2921 | 30850 | 306 | 33947 | 508
BIR1 |  | 0.027 | 0.064516 | 0.076136 | 0.076136 | 13.24 | 14.0115 | 12.8585 | 12.8585 | 0.96 | 0.92529 | 0.91039 | 0.91039 | 0.171 | 0.18004 | 0.171 | 0.171 | 10.2517 | 10.4631 | 10.3723 | 10.3723 | 51.4 | 39.7463 | 53.7576 | 53.7576 | 166 | 147.9867 | 120.0785 | 120.0785 | 126 | 132.2834 | 137.6686 | 137.6686 | 71 | 84.7268 | 70.5323 | 70.5323 | 0.27 | 6.7013 | 0.048907 | 0.048907 | 108 | 114.7416 | 98.7781 | 98.7781 | 16 | 18.2985 | 16.9882 | 16.9882 | 22 | 22.5395 | 22.6096 | 22.6096 | 2 | 2.8711 | 1.5553 | 1.5553 | 30825 | 305 | 32963 | 503
BHVO1 |  | 0.52 | 0.54803 | 0.49491 | 0.49491 | 11.4 | 12.1812 | 11.2242 | 11.2242 | 2.71 | 2.8721 | 2.5931 | 2.5931 | 0.168 | 0.1751 | 0.16971 | 0.16971 | 11.1176 | 11.5353 | 10.8829 | 10.8829 | 45 | 45.1948 | 46.5601 | 46.5601 | 121 | 149.7196 | 127.44 | 127.44 | 136 | 110.3233 | 114.3885 | 114.3885 | 105 | 102.8103 | 94.9698 | 94.9698 | 11 | 14.2096 | 10.1279 | 10.1279 | 403 | 409.4237 | 390.7493 | 390.7493 | 27.6 | 28.4514 | 26.6234 | 26.6234 | 179 | 184.9872 | 170.9685 | 170.9685 | 19 | 22.5145 | 20.557 | 20.557 | 30893 | 303 | 33244 | 497
BHVO1 |  | 0.52 | 0.53355 | 0.48238 | 0.48238 | 11.4 | 11.9258 | 10.9962 | 10.9962 | 2.71 | 2.7612 | 2.4973 | 2.4973 | 0.168 | 0.17369 | 0.16934 | 0.16934 | 11.1176 | 11.3471 | 10.7933 | 10.7933 | 45 | 46.174 | 45.2666 | 45.2666 | 121 | 151.6525 | 135.6513 | 135.6513 | 136 | 113.8382 | 118.1147 | 118.1147 | 105 | 108.8897 | 103.1854 | 103.1854 | 11 | 12.1749 | 7.3965 | 7.3965 | 403 | 399.2913 | 380.7101 | 380.7101 | 27.6 | 28.0366 | 26.2297 | 26.2297 | 179 | 182.1247 | 168.3543 | 168.3543 | 19 | 19.9685 | 18.0942 | 18.0942 | 30939 | 302 | 32699 | 494
BHVO1 |  | 0.52 | 0.52036 | 0.47095 | 0.47095 | 11.4 | 11.8123 | 10.8949 | 10.8949 | 2.71 | 2.7195 | 2.4612 | 2.4612 | 0.168 | 0.16522 | 0.16713 | 0.16713 | 11.1176 | 11.1378 | 10.6936 | 10.6936 | 45 | 44.9974 | 46.8209 | 46.8209 | 121 | 137.5526 | 75.7531 | 75.7531 | 136 | 103.5997 | 107.2608 | 107.2608 | 105 | 108.7176 | 102.9529 | 102.9529 | 11 | 13.601 | 9.3109 | 9.3109 | 403 | 393.7817 | 375.2512 | 375.2512 | 27.6 | 26.9168 | 25.1671 | 25.1671 | 179 | 184.0175 | 170.0829 | 170.0829 | 19 | 18.8772 | 17.0385 | 17.0385 | 31345 | 304 | 32480 | 495
BHVO1 |  | 0.52 | 0.52813 | 0.47767 | 0.47767 | 11.4 | 12.0367 | 11.0952 | 11.0952 | 2.71 | 2.7676 | 2.5028 | 2.5028 | 0.168 | 0.16205 | 0.1663 | 0.1663 | 11.1176 | 11.2729 | 10.7579 | 10.7579 | 45 | 47.4132 | 43.6296 | 43.6296 | 121 | 139.1471 | 82.527 | 82.527 | 136 | 110.7427 | 114.8331 | 114.8331 | 105 | 109.4113 | 103.8903 | 103.8903 | 11 | 14.0561 | 9.9219 | 9.9219 | 403 | 402.3576 | 383.7482 | 383.7482 | 27.6 | 28.9142 | 27.0625 | 27.0625 | 179 | 178.4818 | 165.0273 | 165.0273 | 19 | 24.5747 | 22.5499 | 22.5499 | 30587 | 303 | 32111 | 493
BHVO1 |  | 0.52 | 0.47866 | 0.43483 | 0.43483 | 11.4 | 12.0627 | 11.1184 | 11.1184 | 2.71 | 2.7889 | 2.5213 | 2.5213 | 0.168 | 0.16354 | 0.16669 | 0.16669 | 11.1176 | 11.3722 | 10.8052 | 10.8052 | 45 | 45.3922 | 46.2993 | 46.2993 | 121 | 145.8216 | 110.8808 | 110.8808 | 136 | 110.9034 | 115.0034 | 115.0034 | 105 | 111.8706 | 107.2138 | 107.2138 | 11 | 14.9883 | 11.1732 | 11.1732 | 403 | 410.9454 | 392.257 | 392.257 | 27.6 | 32.2942 | 30.2702 | 30.2702 | 179 | 183.2109 | 169.3463 | 169.3463 | 19 | 16.8549 | 15.0823 | 15.0823 | 31059 | 303 | 32985 | 498
BCR1 |  | 1.69 | 1.9367 | 1.6977 | 1.6977 | 6.95 | 7.6444 | 7.1735 | 7.1735 | 2.24 | 2.6934 | 2.4387 | 2.4387 | 0.18 | 0.21143 | 0.17922 | 0.17922 | 12.1105 | 14.2258 | 12.1642 | 12.1642 | 37 | 51.3221 | 38.466 | 38.466 | 13 | 132.0688 | 52.457 | 52.457 | 19 | 28.0619 | 27.1825 | 27.1825 | 129.5 | 133.4788 | 136.4144 | 136.4144 | 47.2 | 42.8637 | 48.5925 | 48.5925 | 330 | 372.9411 | 354.6023 | 354.6023 | 38 | 39.0231 | 36.656 | 36.656 | 190 | 212.7519 | 196.3252 | 196.3252 | 14 | 13.8569 | 12.1822 | 12.1822 | 31240 | 310 | 34063 | 510
BCR1 |  | 1.69 | 1.9004 | 1.6662 | 1.6662 | 6.95 | 7.5357 | 7.0764 | 7.0764 | 2.24 | 2.6468 | 2.3984 | 2.3984 | 0.18 | 0.20396 | 0.17726 | 0.17726 | 12.1105 | 14.006 | 12.0595 | 12.0595 | 37 | 50.7889 | 39.1702 | 39.1702 | 13 | 123.4584 | 15.8791 | 15.8791 | 19 | 23.7948 | 22.6589 | 22.6589 | 129.5 | 124.8936 | 124.8127 | 124.8127 | 47.2 | 41.6938 | 47.0221 | 47.0221 | 330 | 365.305 | 347.0365 | 347.0365 | 38 | 39.1335 | 36.7608 | 36.7608 | 190 | 210.9818 | 194.7087 | 194.7087 | 14 | 18.6034 | 16.7736 | 16.7736 | 31957 | 310 | 34518 | 507
BCR1 |  | 1.69 | 1.9258 | 1.6883 | 1.6883 | 6.95 | 7.8586 | 7.3647 | 7.3647 | 2.24 | 2.7947 | 2.5263 | 2.5263 | 0.18 | 0.21654 | 0.18055 | 0.18055 | 12.1105 | 14.5498 | 12.3185 | 12.3185 | 37 | 55.153 | 33.4053 | 33.4053 | 13 | 128.9364 | 39.1505 | 39.1505 | 19 | 22.3357 | 21.1121 | 21.1121 | 129.5 | 134.9674 | 138.426 | 138.426 | 47.2 | 43.4188 | 49.3376 | 49.3376 | 330 | 376.6172 | 358.2446 | 358.2446 | 38 | 39.2564 | 36.8775 | 36.8775 | 190 | 225.9546 | 208.3829 | 208.3829 | 14 | 15.2376 | 13.5178 | 13.5178 | 31618 | 309 | 35468 | 505
BCR1 |  | 1.69 | 2.075 | 1.8175 | 1.8175 | 6.95 | 7.8202 | 7.3304 | 7.3304 | 2.24 | 2.798 | 2.5291 | 2.5291 | 0.18 | 0.20732 | 0.17814 | 0.17814 | 12.1105 | 14.4919 | 12.2909 | 12.2909 | 37 | 52.2052 | 37.2993 | 37.2993 | 13 | 127.4309 | 32.7547 | 32.7547 | 19 | 24.6639 | 23.5803 | 23.5803 | 129.5 | 126.4965 | 126.9787 | 126.9787 | 47.2 | 42.6579 | 48.3162 | 48.3162 | 330 | 382.2637 | 363.8392 | 363.8392 | 38 | 46.7079 | 43.949 | 43.949 | 190 | 221.6109 | 204.416 | 204.416 | 14 | 15.3569 | 13.6332 | 13.6332 | 31260 | 309 | 34831 | 507
BCR1 |  | 1.69 | 1.8535 | 1.6256 | 1.6256 | 6.95 | 6.9958 | 6.5943 | 6.5943 | 2.24 | 2.5307 | 2.2981 | 2.2981 | 0.18 | 0.19344 | 0.17451 | 0.17451 | 12.1105 | 13.275 | 11.7114 | 11.7114 | 37 | 48.3096 | 42.4455 | 42.4455 | 13 | 121.8887 | 9.2108 | 9.2108 | 19 | 30.3794 | 29.6393 | 29.6393 | 129.5 | 121.7146 | 120.5166 | 120.5166 | 47.2 | 38.6794 | 42.9757 | 42.9757 | 330 | 345.0057 | 326.9239 | 326.9239 | 38 | 34.8581 | 32.7034 | 32.7034 | 190 | 202.4672 | 186.9325 | 186.9325 | 14 | 11.5982 | 9.9973 | 9.9973 | 31195 | 306 | 32897 | 499
|  | K2O | K2O calc | K2O Slope Corr | TRAD | CaO | CaO calc | CaO Slope Corr | TRAD | TiO2 | TiO2 calc | TiO2 Slope Corr | TRAD | MnO | MnO calc | MnO Slope Corr | TRAD | FeO | FeO calc | FeO Slope Corr | TRAD | Co ppm | Co calc | Co slope |  | Ni ppm | Ni calc | Ni slope | TRAD | Cu ppm | Cu calc | Cu slope | TRAD | Zn ppm | Zn calc | Zn slope |  | Rb ppm | Rb calc | Rb slope | TRAD | Sr ppm | Sr calc | Sr slope | TRAD | Y ppm | Y calcc | Y Slope | TRAD | Zr ppm | Zr calc | Zr slope | TRAD | Nb ppm | Nb calc | Nb slope | TRAD
BIR1 | Average | 0.027 | 0.052589 | 0.065806 | 0.065806 | 13.24 | 14.6363 | 13.4164 | 13.4164 | 0.96 | 0.97892 | 0.95674 | 0.95674 | 0.171 | 0.18855 | 0.17323 | 0.17323 | 10.2517 | 10.9081 | 10.5842 | 10.5842 | 51.4 | 42.9485 | 49.5275 | 49.5275 | 166 | 158.2455 | 163.6591 | 163.6591 | 126 | 136.6006 | 142.2453 | 142.2453 | 71 | 87.0175 | 73.6279 | 73.6279 | 0.27 | 7.8829 | 1.6351 | 1.6351 | 108 | 122.3611 | 106.3275 | 106.3275 | 16 | 18.4586 | 17.1401 | 17.1401 | 22 | 24.1617 | 24.0911 | 24.0911 | 2 | 4.4833 | 3.1148 | 3.1148
BHVO1 | Average | 0.52 | 0.52175 | 0.47215 | 0.47215 | 11.4 | 12.0037 | 11.0658 | 11.0658 | 2.71 | 2.7819 | 2.5152 | 2.5152 | 0.168 | 0.16792 | 0.16783 | 0.16783 | 11.1176 | 11.333 | 10.7866 | 10.7866 | 45 | 45.8343 | 45.7153 | 45.7153 | 121 | 144.7787 | 106.4504 | 106.4504 | 136 | 109.8815 | 113.9201 | 113.9201 | 105 | 108.3399 | 102.4424 | 102.4424 | 11 | 13.806 | 9.5861 | 9.5861 | 403 | 403.1599 | 384.5432 | 384.5432 | 27.6 | 28.9226 | 27.0706 | 27.0706 | 179 | 182.5644 | 168.7559 | 168.7559 | 19 | 20.558 | 18.6644 | 18.6644
BCR1 | Average | 1.69 | 1.9383 | 1.699 | 1.699 | 6.95 | 7.5709 | 7.1079 | 7.1079 | 2.24 | 2.6927 | 2.4381 | 2.4381 | 0.18 | 0.20654 | 0.17794 | 0.17794 | 12.1105 | 14.1097 | 12.1089 | 12.1089 | 37 | 51.5558 | 38.1573 | 38.1573 | 13 | 126.7566 | 29.8904 | 29.8904 | 19 | 25.8471 | 24.8346 | 24.8346 | 129.5 | 128.3102 | 129.4297 | 129.4297 | 47.2 | 41.8627 | 47.2488 | 47.2488 | 330 | 368.4265 | 350.1293 | 350.1293 | 38 | 39.7958 | 37.3894 | 37.3894 | 190 | 214.7533 | 198.1531 | 198.1531 | 14 | 14.9306 | 13.2208 | 13.2208

## STable3C-50 KV XRF2
USGS Standard | ID | CaO | Counts /Rh | Ca-Ka Area | Ca-Ka AreaStd | FeO | Counts /Rh | Fe-Ka Area | Fe-Ka AreaStd | Ni ppm | Counts /Rh | Ni-Ka Area | Ni-Ka AreaStd | Rb ppm | Counts /Rh | Rb-Ka Area | Rb-Ka AreaStd | Sr ppm | Counts /Rh | Sr-Ka Area | Sr-Ka AreaStd | Y ppm | Counts /Rh | Y -Ka Area | Y -Ka AreaStd | Zr ppm | Counts /Rh | Zr-Ka Area | Zr-Ka AreaStd | Nb ppm | Counts /Rh | Nb-Ka Area | Nb-Ka AreaStd | Ba ppm | Counts /Rh | Ba-Ka Area | Ba-Ka AreaStd | Rayleigh/ Compoton | Rh-coh/ Rh-Inc
BIR1 | BIR1 | 13.24 | 2919.0757 | 2683 | 89 | 10.2517 | 25517.7047 | 23454 | 155 | 166 | 3836.2508 | 3526 | 69 | 0.27 | 58.7514 | 54 | 31 | 108 | 850.8078 | 782 | 39 | 16 | 157.7585 | 145 | 30 | 22 | 479.8033 | 441 | 36 | 2 | 50.0475 | 46 | 28 | 7.7 | 810.5521 | 745 | 135 | 1.088 | 0.91913
BIR1 | BIR1 | 13.24 | 3148.6815 | 2770 | 90 | 10.2517 | 26627.3876 | 23425 | 155 | 166 | 4039.8606 | 3554 | 70 | 0.27 | 109.124 | 96 | 31 | 108 | 866.1716 | 762 | 39 | 16 | 148.9088 | 131 | 29 | 22 | 439.906 | 387 | 35 | 2 | 55.6987 | 49 | 29 | 7.7 | 602.4553 | 530 | 136 | 1.1367 | 0.87973
BIR1 | BIR1 | 13.24 | 2818.9745 | 2565 | 90 | 10.2517 | 25498.2565 | 23201 | 153 | 166 | 3772.9199 | 3433 | 68 | 0.27 | 134.0799 | 122 | 32 | 108 | 904.4897 | 823 | 39 | 16 | 201.1198 | 183 | 29 | 22 | 465.9825 | 424 | 35 | 2 | 59.3468 | 54 | 28 | 7.7 | 631.9339 | 575 | 135 | 1.099 | 0.90991
BIR1 | BIR1 | 13.24 | 2980.9537 | 2709 | 89 | 10.2517 | 25801.9208 | 23448 | 154 | 166 | 3840.3575 | 3490 | 69 | 0.27 | 113.3401 | 103 | 31 | 108 | 933.1299 | 848 | 39 | 16 | 148.5525 | 135 | 28 | 22 | 423.6498 | 385 | 34 | 2 | 69.3245 | 63 | 29 | 7.7 | 522.6848 | 475 | 137 | 1.1004 | 0.90877
BIR1 | BIR1 | 13.24 | 2774.9873 | 2595 | 89 | 10.2517 | 24826.2452 | 23216 | 154 | 166 | 3805.8497 | 3559 | 68 | 0.27 | 172.1668 | 161 | 31 | 108 | 943.1749 | 882 | 39 | 16 | 101.5891 | 95 | 28 | 22 | 352.8886 | 330 | 34 | 2 | 13.9017 | 13 | 28 | 7.7 | 609.5348 | 570 | 135 | 1.0694 | 0.93514
BHVO1 | BHVO1 | 11.4 | 2359.8903 | 2193 | 85 | 11.1176 | 27170.4838 | 25249 | 159 | 121 | 3901.9436 | 3626 | 69 | 11 | 137.741 | 128 | 27 | 403 | 2841.9838 | 2641 | 58 | 27.6 | 186.1655 | 173 | 30 | 179 | 1849.8183 | 1719 | 53 | 19 | 235.6662 | 219 | 33 | 139 | 1637.8263 | 1522 | 138 | 1.0761 | 0.92928
BHVO1 | BHVO1 | 11.4 | 2572.4608 | 2434 | 87 | 11.1176 | 26725.4796 | 25287 | 159 | 121 | 3581.7871 | 3389 | 68 | 11 | 112.0299 | 106 | 28 | 403 | 2875.7872 | 2721 | 59 | 27.6 | 256.8233 | 243 | 31 | 179 | 1873.8591 | 1773 | 54 | 19 | 169.1018 | 160 | 32 | 139 | 1674.1076 | 1584 | 137 | 1.0569 | 0.94618
BHVO1 | BHVO1 | 11.4 | 2310.7481 | 2230 | 86 | 11.1176 | 26240.9801 | 25324 | 159 | 121 | 3665.0745 | 3537 | 69 | 11 | 108.802 | 105 | 28 | 403 | 2801.9116 | 2704 | 59 | 27.6 | 180.3005 | 174 | 30 | 179 | 1775.0276 | 1713 | 53 | 19 | 148.178 | 143 | 33 | 139 | 1582.2926 | 1527 | 135 | 1.0362 | 0.96506
BHVO1 | BHVO1 | 11.4 | 2420.8967 | 2306 | 87 | 11.1176 | 26488.1368 | 25231 | 160 | 121 | 3673.338 | 3499 | 69 | 11 | 149.0752 | 142 | 29 | 403 | 2776.7874 | 2645 | 58 | 27.6 | 193.1678 | 184 | 31 | 179 | 1779.4535 | 1695 | 53 | 19 | 120.7299 | 115 | 32 | 139 | 1541.1432 | 1468 | 137 | 1.0498 | 0.95254
BHVO1 | BHVO1 | 11.4 | 2328.9901 | 2193 | 86 | 11.1176 | 26541.7792 | 24992 | 158 | 121 | 3762.705 | 3543 | 68 | 11 | 101.9531 | 96 | 28 | 403 | 2790.9649 | 2628 | 59 | 27.6 | 213.4642 | 201 | 31 | 179 | 1842.5891 | 1735 | 54 | 19 | 196.472 | 185 | 32 | 139 | 1777.8064 | 1674 | 136 | 1.062 | 0.94161
BCR1 | BCR1 | 6.95 | 1806.7347 | 1657 | 84 | 12.1105 | 33609.4082 | 30824 | 175 | 13 | 3854.44 | 3535 | 69 | 47.2 | 369.6337 | 339 | 34 | 330 | 2554.725 | 2343 | 56 | 38 | 305.3022 | 280 | 34 | 190 | 2115.3079 | 1940 | 56 | 14 | 142.8378 | 131 | 32 | 681 | 6800.606 | 6237 | 161 | 1.0904 | 0.91712
BCR1 | BCR1 | 6.95 | 1724.9819 | 1597 | 84 | 12.1105 | 33567.4777 | 31077 | 175 | 13 | 3796.6884 | 3515 | 69 | 47.2 | 375.8883 | 348 | 32 | 330 | 2738.1522 | 2535 | 57 | 38 | 278.6758 | 258 | 33 | 190 | 2015.5393 | 1866 | 54 | 14 | 200.9058 | 186 | 32 | 681 | 7203.4466 | 6669 | 161 | 1.0801 | 0.92581
BCR1 | BCR1 | 6.95 | 1604.1255 | 1430 | 83 | 12.1105 | 34548.1518 | 30798 | 175 | 13 | 3947.4948 | 3519 | 70 | 47.2 | 449.8282 | 401 | 33 | 330 | 2841.4335 | 2533 | 57 | 38 | 288.2939 | 257 | 33 | 190 | 2242.4104 | 1999 | 56 | 14 | 143.5861 | 128 | 32 | 681 | 7392.4385 | 6590 | 160 | 1.1218 | 0.89145
BCR1 | BCR1 | 6.95 | 1600.0421 | 1436 | 83 | 12.1105 | 34734.0615 | 31173 | 176 | 13 | 3919.8803 | 3518 | 70 | 47.2 | 400.0105 | 359 | 33 | 330 | 2706.4779 | 2429 | 57 | 38 | 316.4429 | 284 | 33 | 190 | 2278.6115 | 2045 | 56 | 14 | 166.0211 | 149 | 32 | 681 | 7375.1244 | 6619 | 161 | 1.1142 | 0.89748
BCR1 | BCR1 | 6.95 | 1749.5151 | 1659 | 83 | 12.1105 | 31357.3424 | 29735 | 172 | 13 | 3454.7386 | 3276 | 68 | 47.2 | 341.6774 | 324 | 33 | 330 | 2318.9775 | 2199 | 55 | 38 | 287.8949 | 273 | 34 | 190 | 1964.6453 | 1863 | 54 | 14 | 174.0024 | 165 | 32 | 681 | 6384.3064 | 6054 | 159 | 1.0546 | 0.94826
Method #1: Sensitivity Counts/Unit Concentation
|  | CaO |  |  |  | FeO |  |  |  | Ni ppm |  |  |  | Rb ppm |  |  |  | Sr ppm |  |  |  | Y ppm |  |  |  | Zr ppm |  |  |  | Nb ppm |  |  |  | Ba ppm
|  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD |  | Counts/ Conc | Counts Peak | %StD
BIR1 |  | 201.2387 | 2664.4 | 3.355% |  | 2277.558 | 23348.8 | 0.6604% |  | 21.159 | 3512.4 | 1.959% |  | 397.037 | 107.2 | 29.1% |  | 7.587 | 819.4 | 4.76% |  | 8.6125 | 137.8 | 20.9% |  | 17.8818 | 393.4 | 8.846% |  | 22.5 | 45 | 63.11% |  | 75.1948 | 579 | 23.42%
BHVO1 |  | 199.2281 | 2271.2 | 3.795% |  | 2268.1758 | 25216.6 | 0.6305% |  | 29.081 | 3518.8 | 1.95% |  | 10.4909 | 115.4 | 24.26% |  | 6.6199 | 2667.8 | 2.197% |  | 7.0652 | 195 | 15.69% |  | 9.648 | 1727 | 3.092% |  | 8.6526 | 164.4 | 19.71% |  | 11.1871 | 1555 | 8.785%
BCR1 |  | 223.8561 | 1555.8 | 5.361% |  | 2536.7677 | 30721.4 | 0.5683% |  | 267.1231 | 3472.6 | 1.993% |  | 7.5042 | 354.2 | 9.317% |  | 7.2964 | 2407.8 | 2.342% |  | 7.1158 | 270.4 | 12.35% |  | 10.2242 | 1942.6 | 2.842% |  | 10.8429 | 151.8 | 21.08% |  | 9.4476 | 6433.8 | 2.493%
Average Total Counts /wt% or ppm |  | 208.1076 |  |  |  | 2360.8338 |  |  |  | 105.7877 |  |  |  | 8.9976 |  |  |  | 7.1678 |  |  |  | 7.5978 |  |  |  | 12.5847 |  |  |  | 9.7477 |  |  |  | 31.9431
|  | 13.6756 |  |  |  | 152.4354 |  |  |  | 139.7767 |  |  |  | 2.1119 |  |  |  | 0.49625 |  |  |  | 0.87909 |  |  |  | 4.5965 |  |  |  | 1.5487 |  |  |  | 37.4671
|  | 6.571% |  |  |  | 6.457% |  |  |  | 132.1% |  |  |  | 23.47% |  |  |  | 6.923% |  |  |  | 11.57% |  |  |  | 36.52% |  |  |  | 15.89% |  |  |  | 117.3%
Method #2: |  | CaO |  |  |  | FeO |  |  |  | Ni ppm |  |  |  | Rb ppm |  |  |  | Sr ppm |  |  |  | Y ppm |  |  |  | Zr ppm |  |  |  | Nb ppm |  |  |  | Ba ppm
Slope Correction |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept
|  | 1.0318 | -1.0787 |  |  | 0.45194 | 5.6817 |  |  | 4.4436 | -59.2065 |  |  | 1.3182 | -11.12 |  |  | 1.013 | -18.7519 |  |  | 0.9574 | -0.22358 |  |  | 1.2884 | -19.1343 |  |  | 0.95709 | -1.0701 |  |  | 3.2899 | -43.6904
|  | 0.07503 | 0.86435 |  |  | 0.057289 | 0.69993 |  |  | 13.008 | 466.3942 |  |  | 0.12084 | 3.3057 |  |  | 0.056006 | 17.8806 |  |  | 0.1249 | 3.7316 |  |  | 0.054533 | 7.0946 |  |  | 0.14903 | 2.2072 |  |  | 0.083937 | 11.0139
|  | 0.93568 | 0.71936 |  |  | 0.8272 | 0.3391 |  |  | 0.0088965 | 68.6574 |  |  | 0.90151 | 6.7683 |  |  | 0.96178 | 26.3447 |  |  | 0.81883 | 4.1085 |  |  | 0.97724 | 12.4352 |  |  | 0.76033 | 3.7514 |  |  | 0.99161 | 28.6756
|  | 189.1105 | 13 |  |  | 62.2327 | 13 |  |  | 0.11669 | 13 |  |  | 118.996 | 13 |  |  | 327.1278 | 13 |  |  | 58.7555 | 13 |  |  | 558.18 | 13 |  |  | 41.2423 | 13 |  |  | 1536.208 | 13
|  | 97.8598 | 6.7272 |  |  | 7.1561 | 1.4949 |  |  | 550.0699 | 61279.9301 |  |  | 5451.1403 | 595.5227 |  |  | 227040.7731 | 9022.5603 |  |  | 991.766 | 219.434 |  |  | 86313.1031 | 2010.2302 |  |  | 580.3887 | 182.9447 |  |  | 1263204.563 | 10689.737
Element |  | CaO |  |  |  | FeO |  |  |  | Ni ppm |  |  |  | Rb ppm |  |  |  | Sr ppm |  |  |  | Y ppm |  |  |  | Zr ppm |  |  |  | Nb ppm |  |  |  | Ba ppm
|  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept |  |  | Slope | Intercept
Method #3: Traditional |  | 0.004958 | -1.0787 |  |  | 0.00019143 | 5.6817 |  |  | 0.042005 | -59.2065 |  |  | 0.1465 | -11.12 |  |  | 0.14132 | -18.7519 |  |  | 0.12601 | -0.22358 |  |  | 0.10238 | -19.1343 |  |  | 0.098185 | -1.0701 |  |  | 0.10299 | -43.6904
Rh-tube line Corrected Counts |  | 0.00036054 | 0.86435 |  |  | 2.4266e-05 | 0.69993 |  |  | 0.12296 | 466.3942 |  |  | 0.01343 | 3.3057 |  |  | 0.0078136 | 17.8806 |  |  | 0.016439 | 3.7316 |  |  | 0.0043333 | 7.0946 |  |  | 0.015289 | 2.2072 |  |  | 0.0026277 | 11.0139
|  | 0.93568 | 0.71936 |  |  | 0.8272 | 0.3391 |  |  | 0.0088965 | 68.6574 |  |  | 0.90151 | 6.7683 |  |  | 0.96178 | 26.3447 |  |  | 0.81883 | 4.1085 |  |  | 0.97724 | 12.4352 |  |  | 0.76033 | 3.7514 |  |  | 0.99161 | 28.6756
|  | 189.1105 | 13 |  |  | 62.2327 | 13 |  |  | 0.11669 | 13 |  |  | 118.996 | 13 |  |  | 327.1278 | 13 |  |  | 58.7555 | 13 |  |  | 558.18 | 13 |  |  | 41.2423 | 13 |  |  | 1536.208 | 13
|  | 97.8598 | 6.7272 |  |  | 7.1561 | 1.4949 |  |  | 550.0699 | 61279.9301 |  |  | 5451.1403 | 595.5227 |  |  | 227040.7731 | 9022.5603 |  |  | 991.766 | 219.434 |  |  | 86313.1031 | 2010.2302 |  |  | 580.3887 | 182.9447 |  |  | 1263204.563 | 10689.737
USGS Standard | ID | CaO | CaO calc | CaO slope | Trad | FeO | FeO calc | FeO slope | Trad | Ni ppm | Ni calc | Ni slope | Trad | Rb ppm | Rb calc | Rb slope | Trad | Sr ppm | Sr calc | Sr slope | Trad | Y ppm | Y calc | Y slope | Trad | Zr ppm | Zr calc | Zr slope | Trad | Nb ppm | Nb calc | Nb slope | Trad | Ba ppm | Ba calc | Ba slope | Trad
BIR1 | BIR1 | 13.24 | 14.0268 | 13.3941 | 13.3941 | 10.2517 | 10.8088 | 10.5666 | 10.5666 | 166 | 36.2637 | 101.9334 | 101.9334 | 0.27 | 6.5297 | -2.5129 |  | 108 | 118.6994 | 101.4864 | 101.4864 | 16 | 20.7636 | 19.6555 | 19.6555 | 22 | 38.126 | 29.9867 | 29.9867 | 2 | 5.1343 | 3.8439 | 3.8439 | 7.7 | 25.3748 | 39.7893 | 39.7893
BIR1 | BIR1 | 13.24 | 15.1301 | 14.5325 | 14.5325 | 10.2517 | 11.2788 | 10.779 | 10.779 | 166 | 38.1884 | 110.486 | 110.486 | 0.27 | 12.1282 | 4.8668 | 4.8668 | 108 | 120.8429 | 103.6576 | 103.6576 | 16 | 19.5988 | 18.5404 | 18.5404 | 22 | 34.9556 | 25.9021 | 25.9021 | 2 | 5.714 | 4.3987 | 4.3987 | 7.7 | 18.8602 | 18.3572 | 18.3572
BIR1 | BIR1 | 13.24 | 13.5458 | 12.8978 | 12.8978 | 10.2517 | 10.8005 | 10.5629 | 10.5629 | 166 | 35.665 | 99.2732 | 99.2732 | 0.27 | 14.9018 | 8.5229 | 8.5229 | 108 | 126.1888 | 109.0728 | 109.0728 | 16 | 26.4707 | 25.1195 | 25.1195 | 22 | 37.0277 | 28.5718 | 28.5718 | 2 | 6.0883 | 4.7569 | 4.7569 | 7.7 | 19.7831 | 21.3932 | 21.3932
BIR1 | BIR1 | 13.24 | 14.3241 | 13.7009 | 13.7009 | 10.2517 | 10.9292 | 10.621 | 10.621 | 166 | 36.3025 | 102.1059 | 102.1059 | 0.27 | 12.5967 | 5.4845 | 5.4845 | 108 | 130.1845 | 113.1203 | 113.1203 | 16 | 19.552 | 18.4955 | 18.4955 | 22 | 33.6639 | 24.2379 | 24.2379 | 2 | 7.1119 | 5.7366 | 5.7366 | 7.7 | 16.363 | 10.1415 | 10.1415
BIR1 | BIR1 | 13.24 | 13.3344 | 12.6797 | 12.6797 | 10.2517 | 10.5159 | 10.4342 | 10.4342 | 166 | 35.9763 | 100.6564 | 100.6564 | 0.27 | 19.1348 | 14.1027 | 14.1027 | 108 | 131.5859 | 114.5399 | 114.5399 | 16 | 13.3708 | 12.5776 | 12.5776 | 22 | 28.0411 | 16.9935 | 16.9935 | 2 | 1.4261 | 0.29488 | 0.29488 | 7.7 | 19.0819 | 19.0863 | 19.0863
BHVO1 | BHVO1 | 11.4 | 11.3398 | 10.6217 | 10.6217 | 11.1176 | 11.5089 | 10.883 | 10.883 | 121 | 36.8847 | 104.6928 | 104.6928 | 11 | 15.3087 | 9.0592 | 9.0592 | 403 | 396.4959 | 382.8843 | 382.8843 | 27.6 | 24.5024 | 23.2351 | 23.2351 | 179 | 146.9896 | 170.2454 | 170.2454 | 19 | 24.1765 | 22.0689 | 22.0689 | 139 | 51.2732 | 124.9913 | 124.9913
BHVO1 | BHVO1 | 11.4 | 12.3612 | 11.6756 | 11.6756 | 11.1176 | 11.3204 | 10.7978 | 10.7978 | 121 | 33.8583 | 91.2448 | 91.2448 | 11 | 12.4511 | 5.2925 | 5.2925 | 403 | 401.2119 | 387.6615 | 387.6615 | 27.6 | 33.8022 | 32.1387 | 32.1387 | 179 | 148.8999 | 172.7067 | 172.7067 | 19 | 17.3478 | 15.5332 | 15.5332 | 139 | 52.409 | 128.728 | 128.728
BHVO1 | BHVO1 | 11.4 | 11.1036 | 10.378 | 10.378 | 11.1176 | 11.1151 | 10.7051 | 10.7051 | 121 | 34.6456 | 94.7432 | 94.7432 | 11 | 12.0924 | 4.8196 | 4.8196 | 403 | 390.9053 | 377.2212 | 377.2212 | 27.6 | 23.7305 | 22.496 | 22.496 | 179 | 141.0466 | 162.5885 | 162.5885 | 19 | 15.2013 | 13.4788 | 13.4788 | 139 | 49.5347 | 119.2718 | 119.2718
BHVO1 | BHVO1 | 11.4 | 11.6329 | 10.9241 | 10.9241 | 11.1176 | 11.2198 | 10.7524 | 10.7524 | 121 | 34.7237 | 95.0903 | 95.0903 | 11 | 16.5684 | 10.7197 | 10.7197 | 403 | 387.4001 | 373.6706 | 373.6706 | 27.6 | 25.4241 | 24.1174 | 24.1174 | 179 | 141.3983 | 163.0416 | 163.0416 | 19 | 12.3854 | 10.7838 | 10.7838 | 139 | 48.2464 | 115.0338 | 115.0338
BHVO1 | BHVO1 | 11.4 | 11.1913 | 10.4685 | 10.4685 | 11.1176 | 11.2425 | 10.7626 | 10.7626 | 121 | 35.5685 | 98.8442 | 98.8442 | 11 | 11.3312 | 3.8162 | 3.8162 | 403 | 389.3781 | 375.6742 | 375.6742 | 27.6 | 28.0954 | 26.675 | 26.675 | 179 | 146.4151 | 169.5053 | 169.5053 | 19 | 20.1556 | 18.2206 | 18.2206 | 139 | 55.6553 | 139.408 | 139.408
BCR1 | BCR1 | 6.95 | 8.6817 | 7.8791 | 7.8791 | 12.1105 | 14.2362 | 12.1156 | 12.1156 | 13 | 36.4356 | 102.6974 | 102.6974 | 47.2 | 41.0815 | 43.0318 | 43.0318 | 330 | 356.4193 | 342.2882 | 342.2882 | 38 | 40.1828 | 38.2475 | 38.2475 | 190 | 168.0858 | 197.4256 | 197.4256 | 14 | 14.6534 | 12.9545 | 12.9545 | 681 | 212.8972 | 656.7124 | 656.7124
BCR1 | BCR1 | 6.95 | 8.2889 | 7.4738 | 7.4738 | 12.1105 | 14.2185 | 12.1076 | 12.1076 | 13 | 35.8897 | 100.2716 | 100.2716 | 47.2 | 41.7766 | 43.9482 | 43.9482 | 330 | 382.01 | 368.2106 | 368.2106 | 38 | 36.6783 | 34.8923 | 34.8923 | 190 | 160.158 | 187.2115 | 187.2115 | 14 | 20.6105 | 18.6559 | 18.6559 | 681 | 225.5084 | 698.2014 | 698.2014
BCR1 | BCR1 | 6.95 | 7.7082 | 6.8746 | 6.8746 | 12.1105 | 14.6339 | 12.2953 | 12.2953 | 13 | 37.3153 | 106.6062 | 106.6062 | 47.2 | 49.9944 | 54.7804 | 54.7804 | 330 | 396.4191 | 382.8065 | 382.8065 | 38 | 37.9442 | 36.1043 | 36.1043 | 190 | 178.1856 | 210.438 | 210.438 | 14 | 14.7302 | 13.028 | 13.028 | 681 | 231.4249 | 717.6659 | 717.6659
BCR1 | BCR1 | 6.95 | 7.6885 | 6.8543 | 6.8543 | 12.1105 | 14.7126 | 12.3309 | 12.3309 | 13 | 37.0542 | 105.4462 | 105.4462 | 47.2 | 44.4576 | 47.4821 | 47.4821 | 330 | 377.591 | 363.7343 | 363.7343 | 38 | 41.6491 | 39.6513 | 39.6513 | 190 | 181.0622 | 214.1442 | 214.1442 | 14 | 17.0317 | 15.2308 | 15.2308 | 681 | 230.8829 | 715.8827 | 715.8827
BCR1 | BCR1 | 6.95 | 8.4068 | 7.5954 | 7.5954 | 12.1105 | 13.2823 | 11.6845 | 11.6845 | 13 | 32.6573 | 85.9082 | 85.9082 | 47.2 | 37.9744 | 38.9362 | 38.9362 | 330 | 323.5293 | 308.9717 | 308.9717 | 38 | 37.8917 | 36.054 | 36.054 | 190 | 156.1139 | 182.0011 | 182.0011 | 14 | 17.8505 | 16.0144 | 16.0144 | 681 | 199.8647 | 613.8372 | 613.8372
Standards as Unknowns |  | CaO | CaO calc | CaO slope | TRAD | FeO | FeO calc | FeO slope | TRAD | Ni ppm | Ni calc | Ni slope | TRAD | Rb ppm | Rb calc | Rb slope | TRAD | Sr ppm | Sr calc | Sr slope | TRAD | Y ppm | Y calc | Y slope | TRAD | Zr ppm | Zr calc | Zr slope | TRAD | Nb ppm | Nb calc | Nb slope | TRAD | Ba ppm | Ba calc | Ba slope | TRAD
BIR-1 |  | 13.24 | 14.0722 | 13.441 | 13.441 | 10.2517 | 10.8666 | 10.5927 | 10.5927 | 166 | 36.4792 | 102.891 | 102.891 | 0.27 | 13.0582 | 6.0928 | 8.2442 | 108 | 125.5003 | 108.3754 | 108.3754 | 16 | 19.9512 | 18.8777 | 18.8777 | 22 | 34.3629 | 25.1384 | 25.1384 | 2 | 5.0949 | 3.8062 | 3.8062 | 7.7 | 19.8926 | 21.7535 | 21.7535
BHVO-1 |  | 11.4 | 11.5258 | 10.8136 | 10.8136 | 11.1176 | 11.2813 | 10.7802 | 10.7802 | 121 | 35.1361 | 96.9231 | 96.9231 | 11 | 13.5503 | 6.7415 | 6.7415 | 403 | 393.0783 | 379.4223 | 379.4223 | 27.6 | 27.1109 | 25.7324 | 25.7324 | 179 | 144.9499 | 167.6175 | 167.6175 | 19 | 17.8533 | 16.0171 | 16.0171 | 139 | 51.4237 | 125.4866 | 125.4866
BCR-1 |  | 6.95 | 8.1548 | 7.3354 | 7.3354 | 12.1105 | 14.2167 | 12.1068 | 12.1068 | 13 | 35.8704 | 100.1859 | 100.1859 | 47.2 | 43.0569 | 45.6357 | 45.6357 | 330 | 367.1937 | 353.2022 | 353.2022 | 38 | 38.8692 | 36.9899 | 36.9899 | 190 | 168.7211 | 198.2441 | 198.2441 | 14 | 16.9753 | 15.1767 | 15.1767 | 681 | 220.1156 | 680.4599 | 680.4599

## XRF2 ALL USGS
10 KV | Element | MgO | Al2O3 | SiO2 | P2O5 | K2O | CaO | TiO2 | Cr ppm | MnO | FeO | [10 empty] | Rh-La Coh Area
Total Counts | Avg | 272.7173 | 1786.7687 | 4160.4829 | 5796.5592 | 26368.3136 | 33338.1798 | 33381.6824 | 33.0999 | 39000.4845 | 36950.022 | [10 empty] | 36059.6
/wt% or ppm | std | 60.4495 | 52.4049 | 186.0376 | 446.206 | 2980.5178 | 396.9581 | 2249.7635 | 34.4367 | 2532.4966 | 1359.3157 | [10 empty] | 515.3084
| %std | 22.17% | 2.933% | 4.472% | 7.698% | 11.3% | 1.191% | 6.74% | 104% | 6.494% | 3.679% | [10 empty] | 1.429%
30 KV | Element |  |  |  |  | K2O | CaO | TiO2 |  | MnO | FeO | Co ppm | Ni ppm | Cu ppm | Zn ppm | Rb ppm | Sr ppm | Y ppm | Zr ppm | Nb ppm |  | Rh-Ka Coh Area | Rh-Ka Inc Area
Total Counts | Avg |  |  |  |  | 1127.1006 | 1867.0908 | 3636.9792 |  | 15929.9053 | 24836.0648 | 83.4312 | 29.5979 | 9.4419 | 14.2873 | 75.9578 | 49.0683 | 45.3491 | 57.5492 | 50.2812 |  | 31061.7333 | 33619.6667
/wt% or ppm | std |  |  |  |  | 511.3576 | 9.78 | 325.7222 |  | 830.0131 | 1430.6731 | 21.6357 | 6 | 2.2864 | 1.5173 | 20.2187 | 2.2567 | 2.1176 | 1.978 | 1.7311 |  | 353.8444 | 840.4056
| %std |  |  |  |  | 45.37% | 0.5238% | 8.956% |  | 5.21% | 5.76% | 25.93% | 18.8% | 24.22% | 10.62% | 26.62% | 4.599% | 4.67% | 3.437% | 3.443% |  | 1.139% | 2.5%
50 KV | Element | [5 empty] | CaO |  |  |  | FeO |  | Ni ppm |  |  | Rb ppm | Sr ppm | Y ppm | Zr ppm | Nb ppm | Ba ppm
Total Counts | Avg | [5 empty] | 208.1076 |  |  |  | 2360.8338 |  | 25.12 |  |  | 8.9976 | 7.1678 | 7.5978 | 9.9361 | 9.7477 | 10.3173
/wt% or ppm | std | [5 empty] | 13.6756 |  |  |  | 152.4354 |  | 5.6017 |  |  | 2.1119 | 0.49625 | 0.87909 | 0.40741 | 1.5487 | 1.23
| %std | [5 empty] | 6.571% |  |  |  | 6.457% |  | 22.3% |  |  | 23.47% | 6.923% | 11.57% | 4.1% | 15.89% | 11.92%

## Avg USGS
Table ##-A. USGS standards treated as unknowns using average counts on each standard. Columns compare USGS Reference Values to Calculated Values (mass absorption corrected), slope-corrected concentrations, and concentrations derived by traditional regression of tube-line normalized count data.
|  | BIR-1 |  |  |  | BHVO-1 |  |  |  | BCR-1
10 KV XRF-1 | Reference value | Calculated | Slope Corrected | Traditional Regressed | Reference value | Calculated | Slope Corrected | Traditional Regressed | Reference value | Calculated | Slope Corrected | Traditional Regressed
SiO2 | 47.77 | 46.3531 | 47.7789 | 47.7789 | 49.94 | 49.7232 | 49.9893 | 49.9893 | 54.11 | 55.9169 | 54.0518 | 54.0518
TiO2 | 0.96 | 0.90746 | 0.94496 | 0.94496 | 2.71 | 2.7051 | 2.6436 | 2.6436 | 2.24 | 2.3642 | 2.3214 | 2.3214
Al2O3 | 15.35 | 15.5973 | 15.2929 | 15.2929 | 13.8 | 13.4816 | 13.652 | 13.652 | 13.65 | 13.7435 | 13.8551 | 13.8551
FeO | 10.2517 | 10.219 | 10.3907 | 10.3907 | 11.1176 | 10.8813 | 10.9282 | 10.9282 | 12.1105 | 12.4 | 12.1607 | 12.1607
MnO | 0.171 | 0.16939 | 0.17171 | 0.17171 | 0.168 | 0.16026 | 0.16856 | 0.16856 | 0.18 | 0.18976 | 0.17873 | 0.17873
MgO | 9.68 | 11.5599 | 9.7441 | 9.7441 | 7.23 | 7.7281 | 7.0983 | 7.0983 | 3.48 | 2.5856 | 3.5475 | 3.5475
CaO | 13.24 | 13.421 | 13.2999 | 13.2999 | 11.4 | 11.3587 | 11.3074 | 11.3074 | 6.95 | 6.8822 | 6.9827 | 6.9827
K2O | 0.027 | -0.0088614 | 0.045 | 0.045 | 0.52 | 0.48334 | 0.49542 | 0.49542 | 1.69 | 1.7959 | 1.6966 | 1.6966
P2O5 | 0.046 | 0.045562 | 0.054126 | 0.054126 | 0.273 | 0.2571 | 0.25287 | 0.25287 | 0.36 | 0.38391 | 0.372 | 0.372
Na2O | 1.82 | 1.82 | 1.82 | 1.82 | 2.31 | 2.31 | 2.31 | 2.31 | 3.27 | 3.27 | 3.27 | 3.27
Cr ppm | 382 | 156.9111 | 394.3851 | 394.3851 | 289 | 114.0168 | 265.5019 | 265.5019 | 16 | 34.6772 | 27.1129 | 27.1129
| 99.3157 | 98.264 | 99.5424 | 99.5424 | 99.4686 | 99.0886 | 98.8457 | 98.8457 | 98.0405 | 99.5319 | 98.4366 | 98.4366
30 KV XRF-1 | Reference value | Calculated | Slope Corrected | Traditional Regressed | Reference value | Calculated | Slope Corrected | Traditional Regressed | Reference value | Calculated | Slope Corrected | Traditional Regressed
K2O | 0.027 | 0.052589 | 0.065806 | 0.065806 | 0.52 | 0.52175 | 0.47215 | 0.47215 | 1.69 | 1.9383 | 1.699 | 1.699
CaO | 13.24 | 14.6363 | 13.4164 | 13.4164 | 11.4 | 12.0037 | 11.0658 | 11.0658 | 6.95 | 7.5709 | 7.1079 | 7.1079
TiO2 | 0.96 | 0.97892 | 0.95674 | 0.95674 | 2.71 | 2.7819 | 2.5152 | 2.5152 | 2.24 | 2.6927 | 2.4381 | 2.4381
MnO | 0.171 | 0.18855 | 0.17323 | 0.17323 | 0.168 | 0.16792 | 0.16783 | 0.16783 | 0.18 | 0.20654 | 0.17794 | 0.17794
FeO | 10.2517 | 10.9081 | 10.5842 | 10.5842 | 11.1176 | 11.333 | 10.7866 | 10.7866 | 12.1105 | 14.1097 | 12.1089 | 12.1089
Co ppm | 51.4 | 42.9485 | 49.5275 | 49.5275 | 45 | 45.8343 | 45.7153 | 45.7153 | 37 | 51.5558 | 38.1573 | 38.1573
Ni ppm | 166 | 158.2455 | 163.6591 | 163.6591 | 121 | 144.7787 | 106.4504 | 106.4504 | 13 | 126.7566 | 29.8904 | 29.8904
Cu ppm | 126 | 136.6006 | 142.2453 | 142.2453 | 136 | 109.8815 | 113.9201 | 113.9201 | 19 | 25.8471 | 24.8346 | 24.8346
Zn ppm | 71 | 87.0175 | 73.6279 | 73.6279 | 105 | 108.3399 | 102.4424 | 102.4424 | 129.5 | 128.3102 | 129.4297 | 129.4297
Rb ppm | 0.27 | 7.8829 | 1.6351 | 1.6351 | 11 | 13.806 | 9.5861 | 9.5861 | 47.2 | 41.8627 | 47.2488 | 47.2488
Sr ppm | 108 | 122.3611 | 106.3275 | 106.3275 | 403 | 403.1599 | 384.5432 | 384.5432 | 330 | 368.4265 | 350.1293 | 350.1293
Y ppm | 16 | 18.4586 | 17.1401 | 17.1401 | 27.6 | 28.9226 | 27.0706 | 27.0706 | 38 | 39.7958 | 37.3894 | 37.3894
Zr ppm | 22 | 24.1617 | 24.0911 | 24.0911 | 179 | 182.5644 | 168.7559 | 168.7559 | 190 | 214.7533 | 198.1531 | 198.1531
Nb ppm | 2 | 4.4833 | 3.1148 | 3.1148 | 19 | 20.558 | 18.6644 | 18.6644 | 14 | 14.9306 | 13.2208 | 13.2208
50 KV XRF-1 | Reference value | Calculated | Slope Corrected | Traditional Regressed | Reference value | Calculated | Slope Corrected | Traditional Regressed | Reference value | Calculated | Slope Corrected | Traditional Regressed
CaO | 13.24 | 14.0722 | 13.441 | 13.441 | 11.4 | 11.5258 | 10.8136 | 10.8136 | 6.95 | 8.1548 | 7.3354 | 7.3354
FeO | 10.2517 | 10.8666 | 10.5927 | 10.5927 | 11.1176 | 11.2813 | 10.7802 | 10.7802 | 12.1105 | 14.2167 | 12.1068 | 12.1068
Ni ppm | 166 | 36.4792 | 102.891 | 102.891 | 121 | 35.1361 | 96.9231 | 96.9231 | 13 | 35.8704 | 100.1859 | 100.1859
Rb ppm | 0.27 | 13.0582 | 6.0928 | 8.2442 | 11 | 13.5503 | 6.7415 | 6.7415 | 47.2 | 43.0569 | 45.6357 | 45.6357
Sr ppm | 108 | 125.5003 | 108.3754 | 108.3754 | 403 | 393.0783 | 379.4223 | 379.4223 | 330 | 367.1937 | 353.2022 | 353.2022
Y ppm | 16 | 19.9512 | 18.8777 | 18.8777 | 27.6 | 27.1109 | 25.7324 | 25.7324 | 38 | 38.8692 | 36.9899 | 36.9899
Zr ppm | 22 | 34.3629 | 25.1384 | 25.1384 | 179 | 144.9499 | 167.6175 | 167.6175 | 190 | 168.7211 | 198.2441 | 198.2441
Nb ppm | 2 | 5.0949 | 3.8062 | 3.8062 | 19 | 17.8533 | 16.0171 | 16.0171 | 14 | 16.9753 | 15.1767 | 15.1767
Ba ppm | 7.7 | 19.8926 | 21.7535 | 21.7535 | 139 | 51.4237 | 125.4866 | 125.4866 | 681 | 220.1156 | 680.4599 | 680.4599